## MFG Issues
Recall Number | Trade Name/Product | Recall Class | Date Posted | Recalling Manufacturer | Reason for Recall
Z-0406-2013 | INTRABEAM Balloon Applicator Set; - INTRABEAM Balloon Applicator Set;4.0 cm;Carl Zeiss Meditec AG; Oberkochen; GermanyINTRABEAM System is intended for radiation therapy treatments. | 2 | Nov-20-2012 | Carl Zeiss Meditec; Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve.
Z-0298-2013 | Dental implant - Product is Full OSSEOTITE Parallel Walled Certain Implant; Model Number IFOA411. The product is sterile.Product provides a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | 2 | Nov-16-2012 | Biomet 3i; LLC | Biomet 3i recalled Full OSSEOTITE 2 Tapered Certain Implant and Full OSSEOTITE Parallel Walled Certain Implant due to a manufacturing condition where the depth of the implant's internal hex is too shallow. This condition may prevent the driver or abutment from fully engaging in the implant.
Z-0071-2013 | Terumo Backplate Replacement CXFVR 1500 Holder - Terumo Backplate Replacement CXFVR 1500 HolderThe Flexible venous Reservoir is mounted on the Backplate/Holder and secures the flexible reservoir by strategically placed mounting arms. Additionally; there are three port supports the user incorporates to position the lines that enter and exit the venous bag. The Backplate/Holder is designed to allow the user to move the port support location to accommodate different tubing circuit designs to permit mounting flexibility. | 2 | Oct-16-2012 | Terumo Cardiovascular Systems Corporation | Backplate reservoir holders were manufactured with an incorrect configuration of connector supports. The devices were manufactured and configured with two 3/8 inch and one ? inch connector supports instead of one 3/8 inch and two ? inch connector supports (See location of supports as # 7 and # 8 on backplate attachment). The backplates were manufactured as described in manufacturing instructions a
Z-0025-2013 | Eon Mini - 3788 (IPG) - The Eon Mini Implantable Pulse Generator (IPG) is a 16-channel; rechargeable; multi-program system that allows the connection of one or two leads with four; eight; and 16 electrode leads. It is powered by a hermetically sealed battery within a titanium case. It uses microelectronic circuitry to generate constant current electrical stimulation. Stimulation programs can be delivered as either single stimulation or MultiStim programs depending on patients need.Product Usage:Eon Mini system is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome; intractable low back and leg pain. | 2 | Oct-09-2012 | Advanced Neuromodulation Systems Inc. | Eon Mini IPGs manufactured in April 2012 could potentially exhibit a sudden; brief surge in stimulation that would be felt by the patient as uncomfortable or painful. SJMs routine testing; monitoring; and analysis of manufacturing processes identified three lots of Eon Mini IPGs that the internal battery has the potential to come into contact with the internal microcontroller board.
Z-2423-2012 | Asahi KASEI; REXEED(TM)- 15 LX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 15 LX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2199-2012 | SenTec; V-Sign"2 Sensors - SenTec; V-Sign"2 Sensors The SenTec Digital Monitoring System (SMDS) is used for continuous and noninvasive real-time monitoring of ventilation (tcPCO2) and oxygenation (SpO2). | 2 | Aug-13-2012 | Sentec AG | SenTec is recalling a number of V-Sign"2 Sensors due to a problem regarding configuration parameter writing during manufacturing.
Z-2008-2012 | DeRoyal Surgical Kits and Trays - Custom surgical kits and trays:GEO-MED TOTAL KNEE PACK PGYBK; REF 89-6664.04GEO-MED TOTAL HIP PACK; REF 89-6664.03surgical kits | 2 | Jul-13-2012 | DeRoyal Industries Inc | Surgical kits and trays contain Stryker Hytrel Togas which were recalled because the clear tape; that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga; was not applied during the manufacturing process.
Z-1897-2012 | Ethicon - Ethicon Surgiflo Hemostatic Matrix Kit with Thrombin Plus Flex TipDistributed by Ethicon Ethicon Inc.; Somerville; NJ 08876 USAManufactured by Ferrosan Medical Devices Ferrosan Medical Devices A/SSydmarken 5; DK-2860 Soeberg; DenmarkMixed with sterile saline or thrombin solution is indicated for surgical procedures (except opthalmic) for hemostasis when control of capillary; venous and arteriolar bleeding by pressure; ligature; and other conventional procedures is ineffective or impractical. | 2 | Jun-27-2012 | Ethicon; Inc. | The lyophilized thrombin component of the Surgiflo Hemostatic Matrix kit with Thrombin has the potential to be non-sterile due to an undetected air leak which occurred during the manufacturing process.
Z-1809-2012 | T5 Hytrel Zipper Toga - T5 Hytrel Zipper Toga. Product Usage:The togas are intended to provide a barrier between the operating room and the members of the surgical team in order to help protect against contamination and/or exposure of infectious body fluids and harmful microorganisms. | 2 | Jun-14-2012 | Stryker Instruments Div. of Stryker Corporation | The clear tape; that is intended to aid in the prevention of patients fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga; was not applied during the manufacturing process. Stryker has received 5 complaints associated with the defective seams which allowed bleed through of bodily fluids.
Z-1792-2012 | Lipofilter 3000 - Lipofilter 3000; model ASP-CAN-2; labeled in part ***MicroAire; 1641 Edlich Drive; Charlottesville; VA 22911; USA***Product Usage:A fat collection canister intended for use during liposuction procedures. | 2 | Jun-09-2012 | MicroAire Surgical Instruments; LLC | Medical device was not manufactured under Good Manufacturing Procedures and a user may be unable to sterilize the device according to procedures.
Z-1457-2012 | Care Cliner - CARE CLINER 6530 (STD)/6540 (X-LG) products are labeled in part: "***CARE CLINER 6530/6540 (w/NYLON CASTERS)***Winco***ISO 9001-2008 Certified***Outstanding Value and Quality in Our Standard Care Cliner***Lower price does not mean lower quality when you buy a Winco Care Clinger. Our standard Care Cliner comes with all the superior features you might find on higher-priced clinical chairs.***KEY FEATURES***Standard head-rest cover***Padded armrests***Dual fold down side-tables***Removable side panels***Trendelenburg Positioning***1.800.237.3377***5516 SW 1st Lane***Ocala; Florida 334474***www.wincomfg.com***". Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala; FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"The Care Cliner is a multi position recliner with a steel frame used primarily in clinics; hospitals and treatment centers for recovery; dialysis; oncology; examination; infusion. | 2 | Apr-25-2012 | Winco Mfg.; LLC | Winco Manufacturing LLC is recalling the Care Cliner; XL Care Cliner; Drop Arm Care Cliner; Elite Care Cliner; and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this
Z-1386-2012 | ACCU-CHEK Aviva Nano Blood Glucose Meter - ACCU-CHEK Aviva Nano Blood Glucose Meter No product labeling; beyond the labels applied to the meter; was provided with the meter as these meters were never intended to be distributed. A representative copy of product labeling that would have been included in a finished device kit is attached.The Accu-Chek Aviva Nano Blood Glucose meter is intended to be used with Accu-Chek Aviva test strips for the quantitative measurement of glucose (sugar) in fresh capillary whole blood samples drawn from the fingertips or palm. The system is intended to be used by a single person and should not be shared. The system is intended for self-testing outside the body (in vitro diagnostic use) by people with diabetes at home as an aid to monitor the effectiveness of diabetes control. The system should not be used for the diagnosis of or screening of diabetes or for neonatal use. Alternative site testing should be done only during steady-state times (when glucose is not changing rapidly). | 2 | Apr-09-2012 | Roche Diagnostics Operations; Inc. | Accu-Chek Aviva Nano Blood Glucose meters that were manufactured for the purpose of validating a manufacturing process for product to be distributed outside the US were distributed in the US through eBay by an unauthorized third party. These meters were not intended for distribution; were not verified to meet final product specifications and were not distributed to customers with the necessary pr
Z-1148-2012 | AdvanSync; an orthodontic expansion screw - AdvanSync is an orthodontic device intended for the orthodontic treatment of Class II malocclusions. The AdvanSync part numbers and lot numbers affected by this recall are as follows:Part Number: 600-4000 (Kits-include 10 Right 600-4000 and 10 Left Assemblies). Lot Numbers: 070990588; 070976869; 060971515; 060971514; 04094278DPart Number: 600-4040 (5-Pack; Left Assembly). Lot Numbers: 080901333; 080902356; 070981416; 060971299; 060976084Part Number: 600-4041 (5-Pack; Right Assembly). Lot Numbers: 080902357; 080901335; 070981417; 060971304; 060976079 | 2 | Mar-05-2012 | Ormco Corporation | Manufacturing Error: There is a potential for a braze joint failure in the upper telescopic rod assembly. Although an incident could occur in which the eyelet loses connection from the rod; there is no risk of swallowing or aspirating the appliance because it will still be attached to the patient's crown.
Z-1147-2012 | Titanium Surfix Alpha Screws - Titanium SURFIX ALPHA SCREWS; a component of the Newdeal Hallu Lock Plate System. Surfix Alpha Variable angle locking system exists in different diameters and lengths. The product can be sold as a replenishment part individually packaged sterile; or as one (1) unit which is kitted within an instrumentation tray at which point; the product is non-sterile until the tray in which it is kitted is sterilized a the designated facility. The product is indicated for use in fixation of fractures; osteotomies or arthrodesis of the first metatarso-phalangeal joint. | 2 | Mar-02-2012 | Integra LifeSciences Corp. | Manufacturing error. The part number which corresponds to the length dimension of : 14 mm. was incorrectly etched on to screws with the actual measurement of: 12 mm.
Z-1136-2012 | Orthos CM Bracket; Orthodontic Metal Bracket; - Brand name: Orthos CM Bracket; Orthodontic Metal Bracket;Part Numbers: 454-1210; 454-0210.The intended use of this device is a fixed attachment on a tooth which holds an archwire during orthodontic treatment. | 3 | Mar-01-2012 | Ormco Corporation | The recall was initiated because Ormco has confirmed that manufacturing error resulted in the Orthos CM Brackets being incorrectly manufactured with a torque of -9?. The packaging was labeled with +9? torque; however the brackets contained inside the packaging are actually Orthos CM Brackets with a torque of -9?.
Z-1057-2012 | Powerheart AED; CardioVive AED; GE Responder AED; Nihon Kohden CardioLife AED - Powerheart G3 Automated External Defibrillators (Non-Wearable) and include the following models: Powerheart 9300A; 9300E; 9300P; 9390A; 9390E; CardioVive 92532; 92533; Nihon-Kohden 9200G; 9231; and GE Responder 2019198; 2023440 AEDs.The Powerheart AED G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation; if the victim is breathing; the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs; the device will charge automatically and be able to deliver therapy. The device is intended to be used by personnel who have been trained it is operation. The user should be qualified by training in basic life support or other physician-authorized medical response. | 1 | Feb-27-2012 | Cardiac Science Corporation | The affected Automated External Defibrillators (AEDs) contains a circuit board manufactured with a component that may fail unexpectedly due to a supplier manufacturing defect. If the component were to fail during a rescue attempt; the AED may not deliver defibrillation therapy.
Z-0803-2012 | Salter Labs Bubble humidifier - Salter Labs Bubble humidifier; with 6 PSI safety valve; 350cc; Latex Free; Single Patient use; used with supplemental oxygen therapy. Model Numbers 7600-0-50; E7600-0-50; NP7600-0-50 and special order patient kits that contain 7600 humidifiers. Patient kits have brand names of Salter Labs; AdvaCare; Allcare Medical; Memorial Home Services; Therapy Support Inc; Mendo-Lake Home Respiratory Services; MedAssurance Inc.; Abundant Home Care.Product Usage: Humidifier - single patient use. used with supplemental oxygen therapy. | 1 | Jan-26-2012 | SALTER LABS | A manufacturing defect was discovered internally which indicates that in a small number of humidifier lids; the pop-off may not occur at the appropriate pressure and can result in a leak.
Z-0512-2012 | Urine Reagent Strips (URS) 10 Parameters - Urine Reagent Strips (URS) 10 ParametersURS-l0-1OOW510(K) number: K970250Devise Listing number: D072377Each device consists of 100 strips of Urine Reagent Strips 10 parameterspacked with a desiccant pouch inside a gray bottle with Teco DiagnosticsLabel on the bottle and the bottle placed inside a Teco Diagnostics Box.For the semi-quantitative and qualitative detection of Glucose; Bilirubin; Ketone; Specific Gravity; Blood; pH; Protein; Urobilinogen; Nitrite; Leukocytes in urine. | 2 | Jan-11-2012 | Teco Diagnostics | This recall was initiated in response to findings during a FDA inspection of the firm. The firm's recall has been initiated due to improper manufacture of sub-component nitrite paper Lot 35134; Expiration Date: 01/2014. Discrepancies were found on the information recorded on formulation sheet and manufacturing log sheets (Refer to next paragraph). Sampling size submitted for QC testing was insuff
Z-0463-2012 | Smith& Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE Resector - Smith & Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE ResectorCatalog Number: 7205962. | 2 | Jan-11-2012 | Smith & Nephew; Inc. Endoscopy Division | Manufacturing issue: The distal tip of the outer sheath (return electrode) may detach and become loose in the joint during use.
Z-0296-2012 | Respironics Trilogy 100; 200 and 202 Ventilators - Trilogy 100; 200 and 202 VentilatorsThe device is intended to be used in home; institution/hospital; and portable applications such as wheelchairs and gurneys; and may be used for both invasive and non-invasive ventilation. | 1 | Jan-11-2012 | Respironics; Inc. | Trilogy 100; 200 and 202 units are being recalled due to low impeller press-force used during the manufacturing process. If the press force were low enough; theoretically the impeller could move up the shaft and come into contact with the inside of the blower housing/impeller enclosure. This may impact therapy.
Z-0295-2012 | GammaMed Flexible Applicator Probe - GammaMed Flexible Applicator Probe; Part GM11002560 - 3.2 mm diameter braced PVDF. Varian Medical Systems; Palo Alto; CA; Mfg. by Varian Medical Systems; Haan GmbH. Designed for brachytherapy intracavity treatments in areas such as cancer of the vagina; vaginal stump or rectum. | 2 | Nov-30-2011 | Varian Medical Systems; Inc. | GammaMed Flexible Applicator Probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave.
Z-0271-2012 | TUM-E-VAC - TUM-E-VAC REF 2075; Rx ONLY; NON STERILE --- Ethox International; Buffalo; NY 14204; USA --- Intended use: gastric lavage | 3 | Nov-22-2011 | Ethox International; Inc. | Ethox Tum-E-Vac Gastric Lavage Kit #2075; Lot Number 031123255 was labeled with expiration date symbol and date of 2010/10 on the unit product label. The carton was labeled with a manufacturing date symbol and date of 2010/10. The Ethox Tum-E-Vac Gastric Lavage Kit #2075 does not have an expiration date.
Z-0270-2012 | Uplift Commode Assist - "***UPLIFT TECHNOLOGIES INC.***Bringing quality to life.***Model No.***CA200***Description***Lifting commode chair***Weight Range***80-300+ lb-136+ kg***"Labeling on product: "***1C0101 and K0061014***".A lifting commode chair that is self-powered (non-electric) and provides lifting assistance for those who have reduced mobility due to arthritis; post-operative conditions; degenerative joint disease; Parkinson's disease or other neuromuscular disorders. | 2 | Nov-22-2011 | Uplift Technologies Inc. | On June 26; 2011 Uplift Technologies Inc. Nova Scotia; Canada initiated a recall of the Uplift Commode Assist Model No. CA200; Lot No. 1C0101. The plastic seat does not meet manufacturing specifications. The seat may not fit the unit properly and may interfere with the unit's normal locking mechanism.
Z-0200-2012 | Spirit MB Brackets - Spirit MB Brackets; part #494-1210.The brand name of the device is Spirit MB Brackets; an orthodontic bracket. Spirit MB Brackets have been assigned the product code 76 DYW (Plastic Orthodontic Bracket) by the FDA and is classified as a Class II Medical Device. Spirit MB Brackets are to be fixed attachment on a tooth which holds an archwire during orthodontic treatment. | 3 | Nov-16-2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error.
Z-0171-2012 | OSSEOTITE Certain Implant - ***REF IOSS511***OSSEOTITE Certain Implant 5 x 11.5mm***LOT 910695***Sterile using Radiation***Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park; Ed; 1; Planta 1 Pl. de la Pau s/n 08940 Cornella de Llobregat (Barcelona) Spain***Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment. The Thread form titanium screws (cylindrical dental implants) designed to be placed into mandibular or maxillary jaws in order to affix a dental prosthesis (edentulous or partially edentulous segments) that replaces 1 or more missing teeth. | 2 | Nov-08-2011 | Biomet 3i; LLC | The product may contain a manufacturing defect; that if present; may prevent the driver; cover screw or abutment from fully engaging the implant.
Z-3279-2011 | Zimmer Trabecular Metal Modular Acetabular System - The Trabecular Metal Modular Acetabular System is a modular acetabular cup system consisting of a shell with a preassembled locking ring; a liner; and optional screws. The Shells are available in a variety of sizes and are designed to be used with Trilogy Acetabular System liners. The shell implants are used to replace the hip socket during total hip arthroplasty.The device is indicated for primary or revision surgery for rehabilitating hips damaged as a result on noninflammatory degenerative joint disease (NIDJD) or its composite diagnoses of osteoarthritis; avascular necrosis;protrusio acetabuli; traumatic arthritis; slipped capital epiphysis; fused hip; fracture of the pelvis and diastrophic variant. The device is intended for either cemented or noncemented use. | 2 | Sep-27-2011 | Zimmer Inc. | Lot # 61132698 and Lot 001122467 (added 9/ 18/2012) may have a manufacturing non conformance that could prevent the devices from functioning as intended. Specifically; devices in the affected lot may have a locking ring groove depth that prevents the corresponding liner from snapping into the shell. The groove provides the locking ring room to displace when inserting a liner into the shell. Th
Z-3177-2011 | Monomid Non-Absorbable Suture - Monomid Non-Absorbable Suture are blue; monofilament synthetic absorbable suture; with different sizes and different needle sizes.The products are labeled in parts.These sutures are intended for use in general soft tissue approximation and/or ligation including use in cardiovascular; ophthalmic; and neurological surgery. | 2 | Sep-09-2011 | C P Medical Inc. | Unapproved dye was used in manufacturing the non-absorbable blue nylon suture material.
Z-3013-2011 | GammaMed Flexible Applicator - GammaMed Flexible Applicator probe; 3.2mm diameter; braced PVDF.Varian Medical Systems; Palo Alto; CA.Mfg by Varian Medical Systems; Haan GmbH.Flexible Applicator Robes are designed for brachytherapy intracavitary treatments in areas such as vagina; vagina stump or rectum. | 2 | Aug-12-2011 | Varian Medical Systems; Inc. Oncology Systems | Certain GammaMed Flexible applicator probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave.
Z-2946-2011 | Vascular Solutions; Pronto¿ V4; Extraction Catheter - Vascular Solutions; Pronto¿ V4; Extraction Catheter; Sterilized using ethylene oxide; Rx Only. The following sizes/model numbers are affected: 5.5F( 4005); 6F(4006; 7F(4007; 8F(4008) are involved in the recall.The Pronto catheter is indicated for the removal of fresh; soft emboli and thrombi from vessels in the coronary and peripheral vasculature. | 2 | Aug-04-2011 | Vascular Solutions; Inc. | Vascular Solutions ; Inc recently discovered that there was a manufacturing defect with Pronto V4 extraction catheters (Models 4005; 4006; 4007; 4008). Use of this defective product could lead to a tip separation. This maycause injury and/or procedural delay; resulting in a deteriorating hemodynamic condition of the patientwhich may require medical intervention.
Z-2838-2011 | Incise Zirconia Coping (CBZR0102) - "***3i Incise Custom-Made Label. QP157 Approved 07/09/10*** 3i inc se Custom-made Dental Framework***Biomet 3i 4555 Riverside Drive Palm Beach Gardens; FL 33410 800-342-5454 www.biomet3i.com Please consult instructions for use and recommendation***For Dental Lab use ONLY Non-sterile/For single Use***"Indicated for use in restorative procedure to affix the final prosthesis to the abutment. Coping and frameworks provide the base structure for final porcelain application as the final step of creating a restoration. The devices are intended to be used to create final restoration for dental reconstruction with either a prepared tooth or dental implant abutment. | 3 | Jul-16-2011 | Biomet 3i; LLC | On December 1; 2010 Biomet 3i recalled their product Incise Zirconia Coping (CBZR0102). The Coping contains a manufacturing defect. Coping is oversized. All affected direct accounts were notified by letter on December 14; 2010.
Z-2618-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit; Part number 65418-01; Terumo Latin America Seminsa Pediatric Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2383-2011 | PTFE Coated 3 cm Flexible Tip Guidewire - Gyrus ACMI; REF 25 BX; PTFE Coated 3 cm Flexible Tip Guidewire; .035" x 150 cm; Sterile EO; Rx Only; Gyrus ACMI; Inc. 136 Turnpike Road;Southborough; MA 01772-2104 USAUreteral guidewires are used to establish access in the urinary system for the passage of urological devices. An introducer is first used to establish entry to various portions of the urinary tract (urethra; bladder; ureter); then a guidewire is threaded through the introducer. Following placement of the guidewire; various urological devices can be introduced into the targeted portion of the urinary tract for various therapeutic and diagnostic procedures. | 2 | May-26-2011 | Gyrus Acmi; Incorporated | Gyrus ACMI; Inc. wants to bring a recall issue to their customers attention concerning one lot of their 25BX-PTFE Coated 3 cm Flexible Tip Guidewire. As a result of a manufacturing error in the production this lot of guidewires; some units of product were loaded in the dispenser backwards. In such cases; the stiff end of the guidewire would be dispensed from the product first and; if undetected;
Z-2054-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 32 MM O.D. 55 MM O.D.; REF 71338950; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06 | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-1966-2011 | Container; IV - ExactaMix EVA Container; 2000mL; REF: 740; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1690-2011 | DePuy CEMENT RESTRICTOR; Size 5 - DePuy; a Johnson & Johnson company; CEMENT RESTRICTOR; Size 5; Sterile; DePuy Orthopaedics; 700 Orthopaedic Drive; Warsaw; IN; 46582; 800-366-8143/This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis | 2 | Mar-17-2011 | Depuy Orthopaedics; Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised; the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field; including the implant. Non-sterile product and/or contaminat
Z-0906-2011 | InfraVision IR Illuminator Console - InfraVision IR Illuminator Console; model number 220-180-521; manufactured by Stryker Endoscopy San Jose; for use with the Stryker InfraVision system.Intended use:Medical device for human use; during surgical procedures to transilluminate anatomical structures during surgical procedures with infrared light. It can be used in either open surgical interventions or during laparoscopic procedures. infrared light which is not visible to the human eye is detected by all of Stryker's infrared sensitive cameras. The InfraVision console outputs the infrared light via disposable fibers hooked to the console that are contained in either the E-kit or the U-kit. | 2 | Jan-18-2011 | Stryker Endoscopy | Product may not adequately visualize IR light emitted from U- and E- kits due to a manufacturing error.
Z-0824-2011 | Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long - Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long; Catalog 00-2490-003-50; Zimmer; Warsaw; IN.A reusable surgical instrument used for the insertion of a Lag Screw used in the Zimmer Natural Nail system. | 2 | Dec-28-2010 | Zimmer Inc. | The firm has determined that two lots of this instrument were inappropriately manufactured by the supplier. The manufacturing process created a step in the lumen which impedes the Lag Screw Retaining Shaft from passing through the lumen without the use of excessive force.
Z-0831-2011 | Manual Wheelchairs with Transit System Option - Manual Wheelchairs with Transit System Option;Model: Quickie LXI (EILXI);Product is manufactured and distributed by Sunrise Medical; Inc. Fresno; CA Intended to provide mobility to persons restricted to a sitting position and are unable to transfer from their wheelchairs when traveling in motor vehicles; the wheelchair must serve as a vehicle seat. | 2 | Dec-28-2010 | Sunrise Medical Inc | The stabilizer bar was inadvertently omitted during the manufacturing process.
Z-0545-2011 | Allura XPER - Allura XPER X-ray system.The intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0482-2011 | Neuroform 3 Microdelivery Stent System - Neuroform 3 Microdelivery Stent System; Distributed and manufactured by: Boston Scientific Corporation; Fremont; CA | 2 | Nov-26-2010 | Boston Scientific Corporation | The product may contain manufacturing aids (mandrels) which should not be present in the final product. If the mandrel remains in place; the catheter cannot be loaded over the required guidewire; so cannot be used.
Z-0308-2011 | Panta Arthrodesis Nail - Panta Support Device Panta¿ Arthrodesis Nail Instrumentation Tray One unit per package.The PANTA(R) Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. | 2 | Nov-02-2010 | Integra LifeSciences Corporation | Inability to properly align the calcaneal screw placement on a consistent basis; due to a manufacturing defect of of the targeting frame (Panta Support Device).
Z-2478-2010 | Fabius GS - Fabius GS Anesthesia Machine; OR/Anesthesia; Draeger Medical; Inc.; 3135 Quarry Road; Telford; PA 18969inhalation anesthesia machine for use in operating; induction and recovery rooms | 2 | Sep-23-2010 | Draeger Medical Systems; Inc. | potential failures to Electri-Cord Manufacturing AC power cord
Z-2482-2010 | ARCHITECT¿ i1000SR" Processing Module - ARCHITECT¿ i1000SR" Processing Module System For In Vitro Diagnostic Use; Abbott Laboratories Abbott Park; ILProduct List Number; Part # or Model: 01L86-01In Vitro Diagnostic use only. The Abbott ARCHITECT System is designed to perform automated: Chemistry tests; utilizing photometry and potentiometric technology; and Immunoassay tests; utilizing CMIA (chemiluminescent microparticle assay) detection technology. | 2 | Sep-23-2010 | Abbott Laboratories; Inc | During manufacturing; two components of the ARCHITECT i1000SR were reversed which could potentially cause falsely elevated patient results due to a sample-reagent and reagent-sample carryover.
Z-2444-2010 | St. Jude Medical; Model 3085; Dual-Chamber External Pulse Generator - St. Jude Medical; Model 3085; Dual-Chamber External Pulse GeneratorThe PACE 203H / Model 3085 external pulse generator / temporary cardiac pacemaker is designed to be used with cardiac pacing lead systems for temporary atrial or ventricular or A-V sequential pacing. The PACE 203H / Model 3085 has applications where such pacing modes are indicated for therapeutic; prophylactic; or diagnostic purposes. | 2 | Sep-21-2010 | Osypka Medical Inc | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e.; not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t
Z-2436-2010 | Pleatman Sac¿ Tissue Removal System - The Gyrus ACMI Pleatman Sac Tissue Removal System facilitates removal of specimens while minimizing contamination of the abdominal cavity or wound tract. The system contains one (1) Sac and one (1) obturator in a sterile; ready -to -use; single- use procedure tray. | 2 | Sep-19-2010 | Gyrus Acmi; Incorporated | Gyrus Medical; Inc. has issued a Urgent Medical Device Recall for the Gyrus ACMI product 004942-903 - Pleatman Sac? Tissue Removal System; sold as lot 186540FC with an expiration date of 06/2015. As a result of a manufacturing error in the production of the trays used in this lot; it is possible that a crack could form that would compromise the package integrity; and possibly the sterility of th
Z-2242-2010 | CIVCO Surgi-Tip Transducer cover - CIVCO Surgi-Tip Transducer Cover; Ref. 610-834; Sterile (20.3 x 244cm) polyethylene telescopically-folded cover with attached (6cm) NeoGuard tip. The latex-free surgi tipped covers are utilized to cover ultrasound transducers and prevent cross-contamination from occurring with equipment and patients. | 2 | Aug-20-2010 | Civco Medical Instruments Inc | Polyethylene material did not have proper treatment applied during manufacturing. Untreated polyethylene material does not maintain a secure seal with the adhesive tape used to secure the tip to the cover.
Z-2212-2010 | DYNASTY ¿ A-CLASS ¿ POLY LINER - DYNASTY¿ A-CLASS¿ POLY LINER; REF DLXP-LD36; 1 EACH; Rx ONLY; STERILE EO; GROUP D; I.D. 36 mm; LIP 15o; LINER STD; Wright Medical Technology; 4677 Airline Road; Arlington; TN 38002. Product is used in total hip arthroplasty. | 3 | Aug-12-2010 | Wright Medical Technology Inc | The inner and outer product packaging was missing the following information: shelf life; translations; manufacturing date; CE marking and also listed the incorrect sterilization method.
Z-2205-2010 | Pulmonetic Systems - Pulmonetic Systems P/N 19189-001 Patient Circuit SPU W/O PEEP 15 FT 5/Pkg.These circuits are intended for use with the MR Conditional LTV 1200 System with the ventilator utilizing an external exhalation valve; as single patient use (SPU) only patient circuits for adult and pediatric patients. The LTV 1200 ventilator is intended to provide continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation. The ventilator is suitable for use in institutional; home and transport settings; and is applicable for adult and pediatric patients weighing at least 5 kg (11 lbs.); who require the following types of ventilatory support: -Positive Pressure Ventilation; delivered invasively (via ET tube) or non-invasively (via mask). -Assist/Control; SIMV; CPAP or NPPV modes of ventilation. -Breath types including Volume; Pressure Control and Pressure Support. | 2 | Aug-10-2010 | CareFusion 203; Inc. | CareFusion is recalling the P/N 19189-001 Patient Circuit (SPU W/O PEEP 15ft) because some of the circuit sense lines may have been reversed during the manufacturing process. If the sense lines have been reversed; immediately upon connection of the patient circuit to the ventilator and turning the ventilator on; the ventilator will begin to autocycle (with increased pressure delivered) with both
Z-1767-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 55 cm Straight Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 55 cm Straight Basic Kit; Catalog Number/REF 10302809; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1681-2010 | Lyric Hearing Aid - InSound Medical; Inc; Lyric in-canal hearing aid. | 2 | May-26-2010 | InSound Medical; Inc. | Manufacturing error could result in electrolyte leakage from the product's battery.
Z-1611-2010 | 4.5MM Bone Tap; Modular - 4.5MM Bone Tap; Modular; Manufactured by: Blackstone Medical; Inc. 1720 Bray Central Drive; McKinney; TX 75069.Intended to created a screw thread in a hole drilled into bone to facilitate placement of a bone screw. | 2 | May-19-2010 | Blackstone Medical; Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread.
Z-1599-2010 | Eon Mini" Implantable Pulse Generator (IPG) - Eon Mini Implantable Pulse Generator (IPG). Advanced Neuromodulation Systems; Inc. Plano; TX.Indicated as an aid in the management of chronic; intractable pain of the trunk and/or limbs; including unilateral or bilateral pain. | 2 | May-13-2010 | Advanced Neuromodulation Systems; Inc | Due to a missed inspection step during manufacturing; a small number of devices may experience a gradual or sudden loss of charging capability.
Z-1559-2010 | VENTAK PRIZM Implantable Cardioverter Defibrillator - VENTAK PRIZM Automatic Implantable Cardioverter Defibrillator (AICD). DR; model 1851; DR HE; model 1853 and 1858; PRIZM VR HE; model 1857; PRIZM 2 VR model 1860 and PRIZM 2 DR model 1861. The ICDs are designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation (VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-energy shocks using either a biphasic or monophasic waveform. Bradycardia pacing; including adaptive-rate features is available to detect and treat bradyarrhythmias and to support the cardiac rhythm after defibrillation therapy. The devices denoted with DR offer dual-chamber bradycardia features (atrial and/or ventricular pacing and sensing); and the devices denoted with VR offer single-chamber bradycardia features (ventricular pacing and sensing). Sterilized with gaseous ethylene oxide. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1221-2010 | Point Scanning Confocal Microscope TCS SP5/TCS SP5 X - Shutter HX SP5; a component of the Leica Point Scanning Confocal Microscope Models TCS SP5 and TCS SP5 X.Laboratory instrument. | 2 | May-06-2010 | Leica Microsystems; Inc. | A manufacturing defect was identified that could potentially result in the laser shutter not closing as intended which could result in the shutter not completely closing disrupting the laser beam safely.
Z-1409-2010 | Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 - Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 | 2 | Apr-16-2010 | Boston Scientific Neuromodulation Corporation | Toshiba has determined that certain Sony battery packs installed in Toshiba portable computers can overheat; posing fire hazard to customers. Based on Toshiba's investigation with Sony Corporation; they concluded that the battery cells in specific manufacturing lots could be affected.
Z-1301-2010 | Spanish 100-120V 515B Pulse Oximeter - Spanish 100-120V 515B Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1235-2010 | CADD-Legacy¿ 1 Model 6400 Ambulatory Infusion Pump - CADD-LEGACY¿ 1 Model 6400; Ambulatory Infusion Pump; REF 21-6400-51; Rx Only; Smiths Medical MD; Inc..; St. Paul; MN 55112. | 2 | Apr-05-2010 | Smiths Medical ASD; Inc. | The CADD-Legacy? 1 Model 6400 Ambulatory Infusion Pump (Serial Number 370969) was shipped without documentation that the required accuracy testing had been performed during the manufacturing process. If the subject infusion pump did not go through the accuracy testing; and if it does not meet accuracy specifications; there is a potential for over or under-delivery.
Z-1075-2010 | Medical Action Industries; dialysis kit - Medical Action Industries; Dialysis Kit; using BD Q-Syte Luer Access Split Septum (BD #385100); Part number: 57977. | 1 | Mar-30-2010 | Medical Action Industries Inc | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-1080-2010 | Acacia IV Extension Set with BD Q-Syte - Acacia Catalog # 385166 Dual "Y" IV Extension Set 50cm Macro Bore; Spin Nut w/BD Q-Syte | 1 | Mar-30-2010 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly; possibly resulting in an air embolism.
Z-1186-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 11mm x 20cm x 135¿; Product Part Number: 1038-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1153-2010 | XACT Carotid Stent System - Xact 10 mm x 20 mm x 136 cm (Part Number 82099-01)The Xact Carotid Stent System; used in conjunction with Emboshield Embolic Protection System is indicated for the improvement of the lumen diameter of carotid arteries in patients considered at high risk for adverse events from carotid endarterectomy who require percutaneous carotid angioplasty and stenting for occlusive artery disease. | 2 | Mar-16-2010 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the XACT Carotid Stent System because the product was distributed prior to approval of a PMA supplement for a manufacturing line move between sites. There are no product quality issues associated with this action.
Z-1041-2010 | Pointe Scientific Liquid Hitachi Calcium Reagent Set - Calcium Reagent Set. Pointe Scientific Liquid Hitachi Calcium Reagent Set; Pointe Scientific; Inc; Canton; MI; Catalog #'s HC902-576 and HC902-1000. Calcium Reagent Set for repackaging; Catalog # HC902-SAM-911. The product is used for the quantitative determination of calcium in serum on Hitachi analyzers. | 2 | Mar-10-2010 | Pointe Scientific; Inc. | Non-compliance with the Good Manufacturing Practice regulations; as the manufacturing process for this product had not been validated.
Z-0919-2010 | BCI 3180 Pulse Oximeter - BCI 3180 Pulse Oximeter uses power cord WW3005 (Electri-Cord Manufacturing Co.; 312 E. Main St.; Westfield; PA 16950). | 2 | Mar-02-2010 | Smiths Medical PM; Inc. | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords; manufactured by Electri-Cord Manufacturing Co.; for use with BCI 3180 Pulse Oximeter(Human); 9200 Advisor Vital Signs Monitor(Human or Veterinary); 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq
Z-0914-2010 | Access Peristaltic Pump - Access Peristaltic Pump component of: UniCel DxI 800 Access Immunoassay Systems; UniCel DxI 600 Access Immunoassay Systems; UniCel DxC 880i SYNCHRON Access Clinical Systems.Intended to be used with the Access Immunoassay System to remove waste from the system. | 2 | Mar-02-2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of a new premature failure of the upper aspirate peristaltic pump in the systems listed above. Premature failure of the pump may negatively affect precision. The failure is limited to those UniCel DXI 800; DXI 600; and DxC 880i systems which have a pump identified with one of the new pump manufacturing numbers (Mfg. Nos.) listed ab
Z-0751-2010 | BD Q-Syte Extension Set; Luer access extension - BD Q-Syte Extension Set; 15cm Std Bore; P.V. 1.14 ml Luer-Lok; REF 385101; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0757-2010 | TRIATHLON TS KNEE SYSTEM - Triathlon Total Stabilizer Femoral Component; Size # 2for Use with TS Components; Sterile; Made in the USAStryker Orthopaedics; Mahwah; NJ 07430.The Triathlon Total Knee System components are intended for use in primary and revision total knee arthroplasty to alleviate pain and restore function. | 2 | Feb-01-2010 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing; possibly preventing the assembly of a stem extension.
Z-0745-2010 | Cardiac Cath Pack Kit - Custom Cardiac Cath Pack Kit; K09-08945; Sterile EO; Merit Medical Systems; Inc.; South Jordan; Utah 84095.Intended for fluid management convenience when catheterization is required. | 2 | Jan-29-2010 | Merit Medical Systems; Inc. | A manufacturing defect may prevent adequate sterilization of the entire device.
Z-0488-2010 | Elements Gutta Percha and Elements RealSeal Cartridges - Elements Gutta Percha and Elements RealSeal Cartridges; Part Number: 972-1002; 972-1003; 972-2001; 972-2002 | 2 | Dec-09-2009 | Sybronendo | The recall was initiated because during the manufacturing process of the affected units; the molding process resulted in lock nuts that did not have the level of acceptability required; and therefore; may not withstand the pressures exhibited by the Elements Obturation Unit with which the cartridges are used. More Specifically; there exists a possibility that if the lock nut were to fail; the cart
Z-0198-2010 | VerSys Hip System Femoral Head - Zimmer VerSys Hip System Femoral Head; 12/14 taper; 28 mm dia.; neck length +0 mm; sterile; Zimmer; Warsaw; IN; REF 8018-28-02.Intended for mating femoral stems equipped with tapered necks. | 2 | Nov-13-2009 | Zimmer Inc. | Manufacturing material may be present on the surface of the implant.
Z-0086-2010 | Pivot BiPolar Head - Ortho Development Corp.; Pivot BiPolar Head; 28mm; Hip Prosthesis; Femoral Head. | 2 | Oct-28-2009 | Ortho Development Corporation | A Pivot BiPolar Head; the femoral head in a hip prosthesis; might not move freely after assembly due a stacking up of manufacturing tolerances.
Z-1814-2009 | Engenex - ConvaTec; Engenex Tunnel Dressing; for use with Engenex Negative Pressure Wound Therapy; (1 x 15 cm)Rx only; Latex Free; SterileHealth Industry Bar Code: H984BWS4464509A.Packaged in boxes of 15 units per box. | 3 | Sep-29-2009 | ConvaTec | Labeling: the actual manufacturing date (2/2009) and actual expiration date (2/2011) are switched.
Z-2227-2009 | 3i Custom - 3i Custom # 5753; OSCM2913; 2.9MMX 3.4MM X 13MM; OSSEOTITE CUSTOM; MICROMINI IMPLANT; LOT # 470258; Sterile. Implant Innovations; Inc. 4555 Riverside Drive; Palm Beach Gardens; FL 33410.Custom temporary cover screws are for use in maintaining the soft tissue during the healing process of a dental restoration. | 2 | Sep-22-2009 | Biomet 3i; LLC | Healing Abutment was not made within required manufacturing specification. The product did not have the relief behind the thread; which does not allow proper seating.
Z-2049-2009 | Penumbra Neuron Select Catheter - Penumbra Neuron 5F Select Catheter; Neuro Deliver/Intracranial Access System Catheter; manufactured by Penumbra; Inc.; Alameda; CA.Intended for the introduction of interventional devices into the peripheral; coronary and neuro vasculature. | 1 | Sep-08-2009 | Penumbra Inc. | Leakage-- Due to a manufacturing error; some products may have a fold in the innermost layer; causing disruption in the layering process and may result in blow holds or exposed wire braids; which may result in further damage.
Z-2046-2009 | Terumo Weighted Flexible Sucker - Terumo Sarns cardiovascular disposable products; Weighted Flexible Sucker; sterile; Terumo Cardiovascular Systems Corp.; Ann Arbor; MI.; Catalog No. 804113. | 2 | Aug-04-2009 | Terumo Cardiovascular Systems Corporation | A loose strand of stainless steel left on the device during manufacturing may result in an unretrieved metal fragment remaining in the patient.
Z-1598-2009 | Fusion Omni-Tome - Fusion Omni-Tome; Rx Only; Wilson Cook Medical; GI Endoscopy; 4900 Bethania Station Road; Winston-Salem; NC 27105. | 3 | Jul-21-2009 | Cook Endoscopy | The wire control port of these devices may be missing due to omission of a manufacturing activity.
Z-1009-2009 | Roche COBAS c 111 Analyzer - Roche COBAS c 111 Analyzer with ISE module option; GMMI No. 04777433001; Roche Diagnostics; Indianapolis; IN.Is an in-vitro diagnostic analyzer capable of performing clinical chemistry; specific protein electrolyte tests. Analytes are measured photometrically or turbidimetrically. The analyzer also has an optional ISE module for measuring sodium; potassium and chloride. | 2 | Mar-24-2009 | Roche Diagnostics Corp. | Due to a manufacturing error at the supplier; not all cobas c 111 pinch valve caps are properly glued to the correct position on the pinch valve shaft. Without the glue; the cap can become loose or misaligned. The effect of these faulty valves is that they may not open or close properly which can disturb the fluid transportation within the ISE module resulting in problems with ISE calibration; con
Z-1033-2009 | InteliFUSE; Drill Bit; 2.00 mm; J-Latch - InteliFUSE; Drill Bit; 2.00 mm; J-Latch; QTY: 1; REF: DB.200.JL; non-sterile; 1515 Poydras st.; Ste 1490; New Orleans; LA 70112; MDCI Arundel House; 1 Liverpool Gardens; Worthing; West Sussex BN11 1SL; United KingdomIntended Use: Creating of holes for implant placement. | 2 | Mar-05-2009 | Intelifuse Inc | Manufacturing specification for drill bit was incorrect which; when used in conjunction with the drill guide; could cause the production of metal shavings.
Z-0992-2009 | Bone Wax - Bone Wax. Ethicon; Inc.Bone wax is available sterile in individual foil envelopes; each containing 2.5 grams; and then packaged in an individually sealed over wrap packet.The device is indicated for use to control bleeding from the bone surface during surgical procedures. | 2 | Mar-04-2009 | Ethicon; Inc. | A hospital reported to their Ethicon sales representative that they had an "unmarked" box of bone was on their shipping dock that they could not receive due to lack of appropriate documentation. Upon investigation; this box was determined to be missing from a shipment between two steps of the manufacturing process. Further checks determined that two other boxes were missing.
Z-0508-2009 | Novel VBR Spinal System - Novel PEEK VBR XS 16-22 MM; S-5; 26MM; Part Numbers : 64753-126. Indicated for use in the thoracolumbar spine (T1 to L5) for partial or total replacement of a collapsed; damaged or unstable vertebral body due to tumor or trauma (i.e.; fracture). Intended for use with supplemental spinal fixation systems. | 2 | Dec-31-2008 | Alphatec Spine; Inc. | Firm confirmed that two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number & lot number. They were incorrectly laser marked with the wrong part numbers and sizes. Two lots were mixed from the following lot numbers: Part Number 64733-120 Lot Number 611884; and Part Number 64753-126 and Lot Number 611881.
Z-0153-2009 | Allez Spine Laguna Pedicle Screw System - Allez Spine Laguna Pedicle Screw System; Catalog Number: IT-FN1003Del Mar Pedicle Screw System; IT-ST1001Vertebrae Rotation Tool;IT-H W1001Hex Wrench; IT-CB1001Coronal Bender-L; IT-CB1002 Coronal Bender-R; IT-PH1001Power Rod Holder; IT-BP1001 Thoracic Ball Handle Probe; Curved; IT-BP1002 Thoracic Ball Handle Probe; Straight; DM-MR0530Monoaxial Screw Non-Winged; Size 5 x 30; DM-MR0535 Monoaxial Screw Non-Winged; Size 5 x 35; DM-MR0540 Monoaxial Screw Non-Winged; Size 5 x 40; DM-MR0545 Monoaxial Screw Non-Winged; Size 5 x 45; DM-MW0530 Monoaxial Screw Winged; Size 5 x 30; DM-MW0535 Monoaxial Screw Winged; SizeS x 35; DM-MW0540 Monoaxial Screw Winged; Size S x 40; DM-MW0545Monoaxial Screw Winged; Size 5 x 45; DM-MR0630 Monoaxial Screw Non-Winged; Size 6 x 30; DM-MR0635 Monoaxial Screw Non-Winged; Size 6 x 35;DM-MR0640 Monoaxial Screw Non-Winged; Size 6 x 40; DM-MR0645 Monoaxial Screw Non-Winged; Size 6 x 45; DM-MR0650 Monoaxial Screw Non-Winged; Size 6 x 50; DM-MW0630 Monoaxial Screw Winged; Size 6 x 30;DM-MW0635 Monoaxial Screw Winged; Size 6 x 35; DM-MW0640Monoaxial Screw Winged; Size 6 x 40; DM-MW0645 Monoaxial Screw Winged; Size 6 x 45; DM-MR0735 Monoaxial Screw Non-Winged; Size 7 x 35DM-MR0740 Monoaxial Screw Non-Winged; Size 7 x 40; DM-MR0745 Monoaxial Screw Non-Winged; Size 7 x 45; RD-ST5541 Straight Rod; 4mm HexIntended to help provide correction; immobilization and stabilization of spinal segments as an adjunct to fusion of the thoracic; lumbar and/or sacral space. | 2 | Oct-28-2008 | Allez Spine; LLC | This action is being taken voluntarily as a result of a Warning Letter issued to AlIez Spine LLC by the Food & Drug Administration. The Company decided it could better address the FDA's concerns regarding Current Good Manufacturing Practice requirements of the Quality System without any product in the market so it could focus all of its efforts on responding to the FDA.
Z-2233-2008 | Hamilton Soft Grip Pipettes; 1 ml adjustable volume - Hamilton Soft Grip Pipettes; 1 ml adjustable volume; model number 55019-40Product is indicated for manual pipetting. | 2 | Sep-17-2008 | Hamilton Co | Manufacturing issue with spring may restrict product to drawing less than 1/2 of its stated measured capacity
Z-1183-2008 | Access Power Supply Assembly Sled - Access Power Supply Assembly SledPart Number: A48196/6845C; Component of: Access and Access 2 Immunoassay Systems; SYNCHRON LXi 725 System; and UniCel DxC 600i System; for in vitro diagnostic use. | 2 | Aug-20-2008 | Beckman Coulter Inc | Electrical grounding failure: During manufacturing; internal testing found one instrument wlth an intermittent failure ofthe electrical grounding. The source of the problem was traced to inadequate soldering.
Z-1191-2008 | Integra NeuroSensor - INTEGRA NeuroSensor Probe; REF NS-P; component of Cerebral Blood Flow and Intracranial Pressure Monitoring Kit; Integra NeuroSciences | 2 | Aug-05-2008 | Integra Neuro Sciences | Incorrect pressure reading: Due to a manufacturing error; the ICP reading could be inaccurate to the extent that they exceed label claims. Specifically; the Insertion Error label claim of +/- 3mmHg could be exceeded by up to -3mmHg; that could result in a maximum Insertion Error of +/- 6mmHg.
Z-1556-2008 | CORDIS S.M.A.R.T.¿ Control" Nitinol Stent System - CORDIS S.M.A.R.T.¿ Control" Nitinol Stent SystemModel Number(s): C10040SL; Manufactured in Mexico.Indicated for improving luminal diameter in patients with symptomatic atherosclerotic disease. | 2 | Jul-31-2008 | Cordis Corporation | Data transcription error. The 10 affected lots of Catalog Number C10040SL were manufactured with a prior version of the guidewire lumen material; which is not validated with the current manufacturing process. This condition may result in a separation of the device. If a separation was to occur; and the separated fragment of lumen material was not retrieved on the guidewire; blockage of the vessel
Z-1173-2008 | Trident - Trident PSL Acetabular Shells; Trident PSL HA Solid Back;40 mm to 73 mm; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1132-2008 | Keeler All Pupil II Indirect Ophthalmoscope - Keeler All Pupil II Indirect Ophthalmoscope; Keeler Instruments; Inc.; Broomall; PA 19008 | 2 | May-22-2008 | Keeler Instruments Inc | Good Manufacturing Practices (GMP) deficiencies may compromise the safety and effectiveness of the device. Firm is on import alert 89-04 and released product.
Z-1311-2008 | Terumo "Custom" Cardiovascular Procedure Kit - Cardiovascular Procedure Kit; P/N 64789-01; Wisconsin Cardiovascular Services Pump and Table Pack; Terumo Cardiovascular; Ashland; MA. (the Pump & table pack contains a weighted flexible sucker; Catalog No. 804113) | 2 | Mar-28-2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient.
Z-0460-2008 | Alaris Pump Module (formerly Medley Pump Module) - Cardinal Health Alaris Pump Module (formerly Medley Pump Module); Model 8100; Cardinal Health; San Diego; CA. (All Alaris Pump modules shipped prior to December 23; 2007 are subject to this recall) | 1 | Jan-04-2008 | Cardinal Health 303 Inc DBA Alaris Products | Inaccurate flow rate; related to misassembled (missing; bent or broken) springs during the manufacturing or servicing of the mechanism assembly
Z-0363-2008 | Ascent POCTS (Posterior Occipital Cervical Thoracic System) - Ascent POCTS (Posterior Occipital Cervical Thoracic System) Part Numbers: 65-33 XX Series (65-3310;-3312; -3314; -3318; -3320-; -3322; -3324; -3326; -3328; and -3330); Multi-Axial Screw; 4.0mm x 26mm;Implant grade Titanium Alloy; Blacksone Medical Inc.; 90 Brookdale Drive; Springfield; MA 01104. | 2 | Dec-11-2007 | Blackstone Medical; Inc. | Manufacturing Defect: The saddle of the screw's body relative to the keyway exceeds the tolerance outlined in the device specifications. The extent to which the tolerance is exceeded varies from part to part; in the extreme the tolerance is exceeded by (+) 50%. As a consequence; there is a potential for set-screw cross-threading during implantation. Cross-threading could in turn lead to a comprom
Z-0116-2008 | Boston Scientific - Boston Scientific Radial Jaw Large Capacity Biopsy Forceps; Sterile; 240 mm. Box of 20U PN Number: M00513321REF: 1332-20 | 2 | Oct-31-2007 | Boston Scientific Corporation | Device lacked manufacturing finishing process resulting in endoscope damage.
Z-1207-2007 | TENS Device - Avazzia Tennant Biomodulator¿; Model: BEST AV-1; P/N #TB04303-103; manufactured by Avazzia; Inc.; Dallas; TX 75240. | 3 | Sep-01-2007 | Avazzia; Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications
Z-1042-2007 | OneTouch Ultra Test strip - OneTouch Ultra Test strip; a component of the following OneTouch Ultra Glucose monitoring systems: a)OneTouch Ultra; b) OneTouch InDuo; c) OneTouch Ultra2; d) OneTouch UltraMini; e) OneTouch UltraSmart. 25 strips per vial. Product comes in packs of 100 and 50. (Part numbers: 020-245-07; 020-244-07; 020-963-01) | 2 | Jul-13-2007 | Lifescan Inc | Damaged Vials/Inaccurate test results: a manufacturing equipment problem on one of the packaging lines had the potential to puncture vials; causing a single hole in the side of the vial below the rim. This result of this defect may produce inaccurately high or low blood glucose test results or error message readings.
Z-0829-2007 | Endura No-React - EnDura No-React Dural Substitute 5 cm x 6 cm | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0692-2007 | BioHorizons 0.035 Hex Driver - BioHorizons 0.035 Hex Driver; REF 300-377; BioHorizons Implant Systems; Inc.; Birmingham; AL 35243; non-sterile; Rx only; REF 300-377. | 3 | Apr-13-2007 | BioHorizons Implant Systems Inc | Premature wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear.
Z-0720-2007 | Urologix Targis - Urologix Microwave Delivery System 3.0 - 5.0. For use with prostatic urethra lengths of 3.0-5.0 cm only. Model TA1121D. | 2 | Apr-05-2007 | Urologix; Inc. | The catheter in the kit matches the description on the kit and pouch label; however; an incorrect label was placed onto the catheter during manufacturing.
Z-0658-2007 | Terumo Perfusion System 8000 - Terumo Perfusion System 8000 Base; 5 pump; 220/240v; Model 8000; Catalog No. 16406. | 2 | Mar-27-2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.)
Z-0574-2007 | Abiomed - AB5000 Console Circulatory Support SystemCatalog Number: 0015-0000 | 2 | Mar-07-2007 | Abiomed; Inc. | Unit may alarm ''Low Flow'' and ''Low Pressure'' due to manufacturing material in the pressure source.
Z-0411-2007 | Rocker Reducer - Rocker Reducer; part of the SiLo Spinal System; REF 8880014; Rx only; Material: Stainless Steel; qty: 1 ea; non-sterile; Medtronic Sofamor Danek; 1800 Pyramid Place Memphis; TN 38132; USA; Manufactured at: Bartlett; Tennessee; USA | 2 | Feb-08-2007 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error; the proximal end may deform causing the protrusions that engage the associated implant to break off of the instrument during use.
Z-0300-2007 | BCI Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor - BCI 3404 Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor. Smiths Medical PM; Inc. This device is designed to provide full featured monitoring capabilities in a tabletop design. The system features an ECG cable interface; an SpO2 probe interface; display of patient and waveform data via an EL panel; power status LED; and the function keypad area consisting of six keys (on/off; waveform/trend; alarm silence; menu/enter; up arrow & down arrow). The monitor has a serial port that is used for data communications to a printer or computer and for analog outputs. | 2 | Dec-19-2006 | Smiths Medical PM; Inc. | An error occurred at the board manufacturing site which may affect the following parameters: - An electrical noise observed on the SpO2 pleth waveform generating and displaying erratic heart rates. % SpO2 reading displays dashes. - The ECG waveform may appear noisy; ECG readings appear accurate. -Respiration waveform may appear noisy and the respiration readings will be displayed as dashes.
Z-0243-2007 | BacT/ALERT - bioMerieux BacT/ALERT SN Anaerobic Culture Bottle-REF 259790 | 2 | Dec-05-2006 | bioMerieux; Inc. | Recovery Compromised-during manufacturing ambient air (including oxygen) was inadvertently introduced into the culture bottles; changing their reduction/oxidation potential thus leading to a compromised recovery of some obligate anaerobes and longer detection times.
Z-0190-2007 | ST360 Distal Thread Reduction Guide Pins - Zimmer Spine ST360 Distal Thread Reduction Guide Pins; Part number: 07.00684.00. Non-Sterile. Zimmer Spine; Inc.; 7375 Bush Lake Road; Minneapolis; MN 55439 U.S.A. The Distal Thread Reduction Guide Pin is a reusable instrument that is part of a set of instruments used to perform reduction surgeries with the ST360 Spinal Fixation System. The ST360 Spinal Fixation System is a temporary implant system used to correct spinal deformity and facilitate the biological process of spinal fusion. This system is intended for posterior use in the thoracic; lumbar and sacral areas of the spine. Implants of this system consist of hooks and/or screws connected to rods and are intended to be removed after solid fusion has occurred. This system includes polyaxial screws of varying diameters and lengths; fixed screws of varying diameters and lengths; rods in varying lengths; hooks in varying designs; fixed and adjustable transverse connectors. | 2 | Nov-16-2006 | Zimmer Spine; Inc. | Zimmer Spine; Inc. is conducting a RECALL of all manufacturing lots of ST360 Distal Thread Reduction Guide Pins due to an issue with the thread timing.
Z-0160-2007 | Fusion OMNI - Fusion OMNI ERCP Catheter; Wilson Cook Medical; GI Endoscopy; 4900 Bethania Station Road; Winston-Salem; NC 27105 | 2 | Nov-08-2006 | Cook Endoscopy | Injection through the flush port of these ERCP catheters may be compromised due to omission of a manufacturing activity.
Z-0155-2007 | Access Point Medical Heavy Duty Rollator with Curved Back - Access Point Medical-Heavy Duty Rollator with Curved Back (rolling walker); Product Code RL12011; with seat; curved backrest; 8'' wheels with locks; maximum capacity 400 lbs.;and wire basket. Manufactured for Access Point Medical; St. Louis; MO. Made in China | 2 | Nov-07-2006 | Access Point Medical LLC | The fork component on the wheel of the Rollator can break due to the use of incorrect manufacturing material and/or molding processes.
Z-1491-06 | COBE Autotransfusion set. - Cobe STAT PAC Autotransfusion Set (made with Electa Wash Set); Cobe Product Codes: ASCBFE22; ASCBE22; AS9E55; AS9E22; AS9E12; AS4E55; AS4E22; AS4E17. Cobe Cardiovascular Inc. | 2 | Sep-09-2006 | Cobe Cardiovascular; Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets; possibly causing particles in the blood.
Z-1406-06 | Dri-tek D7120; BG-OX Cartridges Qualicheck2+ - D7120; BG-OX Cartridge; D7120 REF 946-003; and 946-001; cuvettes for the NPT7 series of blood gas analyzers. | 3 | Aug-26-2006 | Radiometer America Inc | QC Bias-a solution used in the manufacturing of the membrane had a different composition which introduces a bias of 20 mpH.
Z-1254-06 | VIASYS Healthcare Disposable Surface Electrodes - VIASYS Healthcare Disposable Bar Electrodes. Re-order No. 019-435600. Single Use only. Manufactured in the U.S.A. for VIASYS Healthcare; 5225 Verona Road; Bldg 2; Madison; WI USA 53711-4495..Each box contains 10 pouches of Bar Electrodes and each pouch contains 10 electrodes | 3 | Jul-25-2006 | Nicolet Biomedical Div of Viasys Healthcare | The Bar Electrodes do not meet performance specification for electrode impedance. A manufacturing defect was discovered in disposable bar electrodes; part number 019-435600; lot 604664; that has the potential to reduce the amplitude of a motor or sensory nerve action potential. These possible lower amplitudes will be measured only if the bar electrode is used as a bar.
Z-1234-06 | Nexcare High Performance Gauze Pad - 3M Nexcare First Aid High Performance Gauze Pad. Individually wrapped & sterile. catalog # 434-10. 3M Consumer Health Care; St. Paul; MN 55144-1000. Made in USA. UPC 05113166774. | 2 | Jul-13-2006 | 3M Company / Medical Division | After reviewing manufacturing records it has been determined that these lots inadvertently were not sterilized after packaging in individual boxes labeled as sterile product.
Z-0913-06 | Bard TruGuide Coxail Biopsy Needle with Depth Stop and additional Blunt Tip Stylet - Bard TruGuide Coxail Biopsy Needle with Depth Stop and additional Blunt Tip Stylet; Catalog Number C1816B; Device Listing Number E148687.The Bard TruGuide Coaxial Biopsy Needle is a three-part device consisting of an outer cannula with an attached female luer lock hub; and inner stylet with an attached male luer lock hub; and a flexible slip rink depth stop. | 2 | May-31-2006 | Bard Peripheral Vascular Inc | Manufacturing change that has introduced variability in the length of the inner stylet. When the inner stylet length is inserted into the cannula component; the exposed stylet length may be up to 8mm longer than expected.
Z-0890-06 | CONTAK RENEWAL 3 - CONTAK RENEWAL 3 CRT-D (model no. H170; H175; H177; H179). Cardiac Resynchronization Therapy Defibrillator. Sterilized with gaseous ethylene oxide. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. CONTAK RENEWAL 3 devices provide ventricular tachyarrhythmia and cardiac resynchronization therapies. | 3 | May-23-2006 | Guidant Corporation | Devices manufactured with batteries built on April 26 or 27; 2005; by a single operator; may have battery welds that may have been performed at settings outside of typical manufacturing specifications at the supplier.
Z-0830-06 | Zinetics Reference pH Catheter - Medtronic Zinetics 24ME Multi-Use External Reference pH Catheter. For Gastroesophageal pH Measurements. Catalog number 9012P2121. Distributed by Medtronic; Inc.; 710 Medtronic Parkway NE; Minneapolis; MN 55432-5604; USA. | 2 | May-03-2006 | Medtronic Gastroenterology / Urology | Lot 0001491 has a manufacturing defect in the sensor spacing. Sensors have been placed 5 cm apart instead of 15 cm apart. Medtronic has determined that patient safety is not compromised. Clinicians however should review and reassess the results of any patient studies for which catheters from lot 0001491 were used.
Z-0743-06 | METRx II Tube - METRx II Tube; 20mm x 6cm stainless steel tubular retractor; Part No. 9569723 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0695-06 | - Fisher & Paykel Healthcare Adult Respiratory Ventilator Circuit-Dual Heated Catalogue Number RT111 | 2 | Mar-29-2006 | Fisher & Paykel Healthcare Inc | Risk of fire. The manufacturing process may have the potential to damage the insulation of the heater wire in circuits within the specified lot range. This may lead to a potential risk of fire.
Z-0612-06 | Baxter Auto Syringe AS50 Infusion Pump - Baxter Auto Syringe AS50 Infusion Pump; product code 1M8550; Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A. | 2 | Mar-16-2006 | Baxter Healthcare Corp. | A non-conforming gasket was used in manufacturing of the specific pumps; causing the pumps to be more susceptible to electro-static discharge; which could increase the potential for the pump to fail during use and alarm.
Z-0504-06 | Hospira Latex Free Blood Secondary Set - Hospira Latex Free Blood Secondary Set; 26 inch with Secure Lock; Made in Costa Rica; Hospira; Inc.; Lake Forest; IL 60045 U.S.A. | 2 | Feb-08-2006 | Hospira Inc. | The I.V. administrations sets were stolen in-transit between the manufacturing site and the sterilizer; resulting in non-sterilized product possibly being on the market.
Z-0392-06 | Medtronic Slimline Catheter - Medtronic Slimline Single-Use Internal Reference pH Catheter for Gastroesophageal pH measurements. Models 9012P3101; 9012P3111; 9012P3121; 9012P3131 and 9012P3141. Distributed by / Medtronic Inc. 710 Medtronic Parkway NE; Minneapolis; MN 55432-5604 USA | 3 | Jan-13-2006 | Medtronic Gastroenterology / Urology | Certain lots have a manufacturing defect that allows the water pushed through the catheter to travel back into the connector of the Digitrapper pH recorder. Medtronic has determined that patient safety is not compromised; but catheter function may be affected and damage may result to the Digitrapper pH recorder.
Z-0354-06 | Respironics PLV Continuum - The PL V Continuum ventilator is a microprocessor controlled; compressorbased; mechanical ventilator. It is intended to control or assist breathing by delivering room air to the patient. PL V Continuum utilizes an internal compressor to generate compressed air for delivery to the patient. Breath delivery is controlled by software algorithms. The user interface on PL V Continuum has a membrane keypad with indicator Light Emitting Diodes (LED) for the selection and acceptance of patient settings and for the display of alarm conditions. PL V Continuum is capable of providing the following types of ventilatory support:. Positive Pressure Ventilation; delivered either invasively (via endotracheal or tracheostomy tube) or non-invasively (via mask or mouthpiece).. Assist/Control; Spontaneous Intermittent Mandatory Ventilation (SIMV) or Continuous Positive Airway Pressure (CP AP) modes of ventilation.. Volume-Controlled (VC). Available in AlC and SIMV.. Pressure-Controlled (PC). Available in AlC and SIMV.. Pressure Support (PS). Available in SIMV and SPONT. | 2 | Jan-06-2006 | Respironics California Inc | During product testing of the PLV Continuum Ventilator's Safety Valve/Pressure Relief Valve (SV/PRV) at our manufacturing facility; Respironics found that when the PLY Continuum Ventilator is operated in the 'Face Up' orientation; the SV /PRV may stick in the open position; which may result in loss of delivered gas volume to the patient.
Z-0266-06 | Symphony DR and ELA Rhapsody DR - Symphony DR model 2550 and ELA Rhapsody DR model 2510. A dual chamber rate-responsive pacemaker that uses both minute ventilation and accelerometer sensors. Sterilization with Ethylene Oxide Gas. Sorin Group; ELA medical; 98; rue Maurice Amoux; 92120 Montrouge - France. Made in France. | 2 | Dec-02-2005 | ela Medical Llc | Through post-market surveillance; ELA Medical has observed that a no-output condition could occur in a limited number of Symphony or rhapsody pacemakers. The no-output condition could occur due to metal migration caused by a specific manufacturing process.
Z-0144-06 | D-Stat Radial Hemostat Band - D-Stat 2 Dry Hemostatic Bandage; model number 3010. Sterile; latex free single use device. Vascular Solutions; Inc. 6464 Sycamore Court; Minneapolis; MN 55369 USA. Each D-Stat 2 Dry hemostatic bandage (D-Stat 2 Dry) consists of the following components: -One (1) scored lyophilized pad consisting of thrombin; sodium carboxylmethyulcellulose and calcium chloride; -Two (2) Adhesive bandages. | 2 | Nov-16-2005 | Vascular Solutions; Inc. | A defect affecting the seals of the paper pouch holding the adhesive bandage provided with the D-Stat products was noted during manufacturing. The integrity of the seals of the D-stat Dry pads which are packaged separately in a foil pouch is not in question.
Z-1585-05 | Bard Luminexx 3 Biliary Stent and Delivery System - Bard Luminexx 3 Biliary Stent and Delivery System | 2 | Sep-24-2005 | Bard Peripheral Vascular Inc | Angiomed Material Review Board reviewing the manufacturing process observed an increase in the internal scrap rate (failure mode: broken stent struts) of nitinol stents and determined the root cause to be a manufacturing change from polymer to aluminum fixtures.
Z-1113-05 | Healthwear - Roche Diagnostic HealthWear brand Professional System for weight management; Catalog # 04344456001. | 3 | Aug-12-2005 | Roche Diagnostics Corp. | Lack of adequate manufacturing controls.
Z-1085-05 | - Puritan Bennett 840 Series Ventilator Power SupplyModel/Catalog Numbers: 4-076314-00; 4-076314-SP | 2 | Aug-06-2005 | Puritan Bennett | Manufacturing defect at contract manufacturer; employee error; poor training; improper inspection.
Z-1068-05 | Sequoia - Sequoia Oxymatic electronic oxygen conservers. Models OM-302; OM-311 ***and OM-312***. | 2 | Aug-03-2005 | Chad Therapeutics Inc | Shot size below manufacturing specifications. Risk of less than the specified quantity of oxygen delivered per pulse.
Z-0852-05 | Continuous Nerve Block Needle - Arrow Continuous Nerve Block Needle; Latex-Free* Rx only*Sterile*Arrow International Inc.; 2400 Bernville; Road; Reading; PA 19605 USA | 2 | Jun-01-2005 | Arrow International Inc | Possible difficulty in passing the catheter through the needle due to a reduced ID of the needle caused by over crimping of the needle during the manufacturing process.
Z-0434-05 | Guiding Catheter - VISTA BRITE TIP RDC(1) Guiding Catheter. Product Code 67021055; lot X0704002; Precutaneous Catheter. | 3 | Jan-26-2005 | Cordis Corporation | Two complaints have been received and confirmed; regarding a lot; indicating that the distal end had an 'RDC' shape rather than the 'RDC(1)' shape. The RDC shape has a larger curve. The product was manufactured with the incorrect forming wire during the guiding catheter manufacturing process.
Z-0393-05 | - Weck; Endo5; Hem-o-lok¿ ML; Ref 544965; Hem-o-lok ML Endoscopic Clip Applier; Non-Sterile Clean and sterilize before use. Weck Closure Systems¿; TFX Medical Ltd.; Halifax Road; P.O. Box 138; High Wycombe HP123NB U.K. | 3 | Jan-13-2005 | Weck | The flush port on some appliers was occluded with an epoxy resin during manufacturing.
Z-1341-04 | ARCHITECT - ARCHITECT c8000 Processing Module; List Number 01G06-01. Manufactured by Toshiba Corporation; Medical Systems Division; Japan. Distributed by Abbott Laboratories Diagnostic Division; Irving; Texas | 3 | Aug-13-2004 | Abbott Laboratories; Inc | Process change in the manufacturing of the AC/DC driver board used on the analyzer may result in one or more of the solid state relays to fail during use.
Z-0992-04 | ABX Diagnostics and Horiba - ABX Diagnostics Minotrol 16 Hematology Controls; Normal and High; in vials distributed by ABX Diagnostics; Irvine; CA; andHoriba LC-TROL 16 (Hematology Reference Controls); Normal and High; in vials distributed by Horiba Ltd.; Kyoto; Japan | 3 | Jul-20-2004 | R & D Systems;Inc | Clots in hematology controls resulted in low red blood cell counts when used in blood analyzers because control values were outside the expected range due to the clotting which took place after manufacturing.
Z-1066-04 | Visulas 532 s with Visulink 532/U - Surgical powered laser instrument; used for photocoagulation of ocular tissues in the treatment of diseases of the eye. | 1 | Jul-20-2004 | Carl Zeiss Ophthalmic Systems Inc | The coating of laser mirrors in a defined manufacturing lot of the Visulink 532/U devices may partially tear open and loosen from the base plate. The mirror may misdirect the laser beam to an unintended target in or on the eye. Retinal bleeding and burns may result.
Z-0722-04 | - Esprit Ventilator. | 2 | Jul-20-2004 | Respironics California Inc | Valves may stick in the closed position due to manufacturing error.
Z-0536-04 | Four Shooter Saeed Multi-Band Ligator - Four Shooter Saeed Multi-Band Ligator | 3 | Jul-20-2004 | Wilson-Cook Medical Inc | Product label lacks the symbols specifying the meaning of the manufacturing and expiration date.
Z-0506-04 | Gas analyzer - AG-920RA Multi-gas unit analyzer | 2 | Jan-31-2004 | Nihon Kohden America Inc | Manufacturing error results in screw affecting electromagnetic interferences. Also metal lid is not tight enough.
Z-0658-03 | Flexgard; off-the-shelf knee brace - Flexgard; off-the-shelf knee brace; product number 14761106 | 2 | Mar-19-2003 | De Royal Industries | Manufacturing defect possibly causing screws to come loose causing the brace to collapse during use.
Z-0630-03 | OneTouch SureStep - OneTouch SureStep Glucose Test Strips (consumer use); OneTouch SureStep Pro Glucose Test Strips (hospital use) | 2 | Mar-12-2003 | Lifescan Inc | The Test Strips have manufacturing errors; in that the 'Confirmation Dot' is off-center or split in half; and defective bottle Cap.

## Recalls
Recall Number | Trade Name/Product | Recall Class | Date Posted | Recalling Manufacturer | Reason for Recall
Z-0406-2013 | INTRABEAM Balloon Applicator Set; - INTRABEAM Balloon Applicator Set;4.0 cm;Carl Zeiss Meditec AG; Oberkochen; GermanyINTRABEAM System is intended for radiation therapy treatments. | 2 | Nov-20-2012 | Carl Zeiss Meditec; Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve.
Z-0404-2013 | INTRABEAM Balloon Applicator Set - INTRABEAM Balloon Applicator Set;3.0 cm;Carl Zeiss Meditec AG; Oberkochen; GermanyINTRABEAM System is intended for radiation therapy treatments | 2 | Nov-20-2012 | Carl Zeiss Meditec; Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve.
Z-0405-2013 | INTRABEAM Balloon Applicator Set - INTRABEAM Balloon Applicator Set;3.5 cm;Carl Zeiss Meditec AG; Oberkochen; GermanyINTRABEAM System is intended for radiation therapy treatments. | 2 | Nov-20-2012 | Carl Zeiss Meditec; Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve.
Z-0298-2013 | Dental implant - Product is Full OSSEOTITE Parallel Walled Certain Implant; Model Number IFOA411. The product is sterile.Product provides a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | 2 | Nov-16-2012 | Biomet 3i; LLC | Biomet 3i recalled Full OSSEOTITE 2 Tapered Certain Implant and Full OSSEOTITE Parallel Walled Certain Implant due to a manufacturing condition where the depth of the implant's internal hex is too shallow. This condition may prevent the driver or abutment from fully engaging in the implant.
Z-0297-2013 | Dental implant - Product is Full OSSEOTITE 2 Tapered Certain Implant; Model Number XIFNT413; and size 4 x 13 mm. The product is sterile.Product provides a means for prosthetic attachment in single tooth restorations and in partially or fully edentulous spans with multiple single teeth utilizing delayed or immediate loading or with a terminal or intermediary abutment for fixed or removable bridgework and to retain over dentures. | 2 | Nov-16-2012 | Biomet 3i; LLC | Biomet 3i recalled Full OSSEOTITE 2 Tapered Certain Implant and Full OSSEOTITE Parallel Walled Certain Implant due to a manufacturing condition where the depth of the implant's internal hex is too shallow. This condition may prevent the driver or abutment from fully engaging in the implant.
Z-0071-2013 | Terumo Backplate Replacement CXFVR 1500 Holder - Terumo Backplate Replacement CXFVR 1500 HolderThe Flexible venous Reservoir is mounted on the Backplate/Holder and secures the flexible reservoir by strategically placed mounting arms. Additionally; there are three port supports the user incorporates to position the lines that enter and exit the venous bag. The Backplate/Holder is designed to allow the user to move the port support location to accommodate different tubing circuit designs to permit mounting flexibility. | 2 | Oct-16-2012 | Terumo Cardiovascular Systems Corporation | Backplate reservoir holders were manufactured with an incorrect configuration of connector supports. The devices were manufactured and configured with two 3/8 inch and one ? inch connector supports instead of one 3/8 inch and two ? inch connector supports (See location of supports as # 7 and # 8 on backplate attachment). The backplates were manufactured as described in manufacturing instructions a
Z-0025-2013 | Eon Mini - 3788 (IPG) - The Eon Mini Implantable Pulse Generator (IPG) is a 16-channel; rechargeable; multi-program system that allows the connection of one or two leads with four; eight; and 16 electrode leads. It is powered by a hermetically sealed battery within a titanium case. It uses microelectronic circuitry to generate constant current electrical stimulation. Stimulation programs can be delivered as either single stimulation or MultiStim programs depending on patients need.Product Usage:Eon Mini system is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs including unilateral or bilateral pain associated with the following: failed back surgery syndrome; intractable low back and leg pain. | 2 | Oct-09-2012 | Advanced Neuromodulation Systems Inc. | Eon Mini IPGs manufactured in April 2012 could potentially exhibit a sudden; brief surge in stimulation that would be felt by the patient as uncomfortable or painful. SJMs routine testing; monitoring; and analysis of manufacturing processes identified three lots of Eon Mini IPGs that the internal battery has the potential to come into contact with the internal microcontroller board.
Z-2423-2012 | Asahi KASEI; REXEED(TM)- 15 LX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 15 LX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2424-2012 | Asahi KASEI; REXEED(TM)- 18 LX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 18 LX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2418-2012 | Asahi KASEI; REXEED(TM)- 15 SX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 15 SX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2420-2012 | Asahi KASEI; REXEED(TM)- 21 SX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 21 SX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2421-2012 | Asahi KASEI; REXEED(TM)- 25 SX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 25 SX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2422-2012 | Asahi KASEI; REXEED(TM)- 13 LX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 13 LX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2419-2012 | Asahi KASEI; REXEED(TM)- 18 SX; HEMODIALYZER - Asahi KASEI; REXEED(TM)- 18 SX; HEMODIALYZERdialysis filter | 2 | Sep-20-2012 | Asahi Kasei Medical Co.; Ltd.; Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products; though there is no sufficient information to date to conclusively link these adverse events to the products.
Z-2199-2012 | SenTec; V-Sign"2 Sensors - SenTec; V-Sign"2 Sensors The SenTec Digital Monitoring System (SMDS) is used for continuous and noninvasive real-time monitoring of ventilation (tcPCO2) and oxygenation (SpO2). | 2 | Aug-13-2012 | Sentec AG | SenTec is recalling a number of V-Sign"2 Sensors due to a problem regarding configuration parameter writing during manufacturing.
Z-2008-2012 | DeRoyal Surgical Kits and Trays - Custom surgical kits and trays:GEO-MED TOTAL KNEE PACK PGYBK; REF 89-6664.04GEO-MED TOTAL HIP PACK; REF 89-6664.03surgical kits | 2 | Jul-13-2012 | DeRoyal Industries Inc | Surgical kits and trays contain Stryker Hytrel Togas which were recalled because the clear tape; that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga; was not applied during the manufacturing process.
Z-1897-2012 | Ethicon - Ethicon Surgiflo Hemostatic Matrix Kit with Thrombin Plus Flex TipDistributed by Ethicon Ethicon Inc.; Somerville; NJ 08876 USAManufactured by Ferrosan Medical Devices Ferrosan Medical Devices A/SSydmarken 5; DK-2860 Soeberg; DenmarkMixed with sterile saline or thrombin solution is indicated for surgical procedures (except opthalmic) for hemostasis when control of capillary; venous and arteriolar bleeding by pressure; ligature; and other conventional procedures is ineffective or impractical. | 2 | Jun-27-2012 | Ethicon; Inc. | The lyophilized thrombin component of the Surgiflo Hemostatic Matrix kit with Thrombin has the potential to be non-sterile due to an undetected air leak which occurred during the manufacturing process.
Z-1809-2012 | T5 Hytrel Zipper Toga - T5 Hytrel Zipper Toga. Product Usage:The togas are intended to provide a barrier between the operating room and the members of the surgical team in order to help protect against contamination and/or exposure of infectious body fluids and harmful microorganisms. | 2 | Jun-14-2012 | Stryker Instruments Div. of Stryker Corporation | The clear tape; that is intended to aid in the prevention of patients fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga; was not applied during the manufacturing process. Stryker has received 5 complaints associated with the defective seams which allowed bleed through of bodily fluids.
Z-1810-2012 | T4 Hytrel Zipper Toga - T4-Hytrel Zipper Toga Product Usage - The togas are intended to provide a barrier between the operating room and the members of the surgical team in order to help protect against contamination and/or exposure of infectious body fluids and harmful microorganisms. | 2 | Jun-14-2012 | Stryker Instruments Div. of Stryker Corporation | The clear tape; that is intended to aid in the prevention of patients fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga; was not applied during the manufacturing process. Stryker has received 5 complaints associated with the defective seams which allowed bleed through of bodily fluids.
Z-1792-2012 | Lipofilter 3000 - Lipofilter 3000; model ASP-CAN-2; labeled in part ***MicroAire; 1641 Edlich Drive; Charlottesville; VA 22911; USA***Product Usage:A fat collection canister intended for use during liposuction procedures. | 2 | Jun-09-2012 | MicroAire Surgical Instruments; LLC | Medical device was not manufactured under Good Manufacturing Procedures and a user may be unable to sterilize the device according to procedures.
Z-1457-2012 | Care Cliner - CARE CLINER 6530 (STD)/6540 (X-LG) products are labeled in part: "***CARE CLINER 6530/6540 (w/NYLON CASTERS)***Winco***ISO 9001-2008 Certified***Outstanding Value and Quality in Our Standard Care Cliner***Lower price does not mean lower quality when you buy a Winco Care Clinger. Our standard Care Cliner comes with all the superior features you might find on higher-priced clinical chairs.***KEY FEATURES***Standard head-rest cover***Padded armrests***Dual fold down side-tables***Removable side panels***Trendelenburg Positioning***1.800.237.3377***5516 SW 1st Lane***Ocala; Florida 334474***www.wincomfg.com***". Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala; FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"The Care Cliner is a multi position recliner with a steel frame used primarily in clinics; hospitals and treatment centers for recovery; dialysis; oncology; examination; infusion. | 2 | Apr-25-2012 | Winco Mfg.; LLC | Winco Manufacturing LLC is recalling the Care Cliner; XL Care Cliner; Drop Arm Care Cliner; Elite Care Cliner; and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this
Z-1461-2012 | Care Cliner - NOCTURNAL ELITE CARE CLINER 6980(STD)/6990(X-LG) products are labeled in part: "***NOCTURNAL ELITE CARE CLINER 6980/6990***Winco***ISO 9001-2008 Certified***An infinite position recliner; independent leg rest and lay-flat all in one chair!***With lay-flat positioning and an independent leg rest; Winco's Nocturnal Elite Care Cliner provides optimal comfort for use in nocturnal dialysis; oncology; acute care and other patient care areas.***HEAT MASSAGE (Optional)***KEY FEATURES***Left side swing-arm for modified entry and ease of cleaning***User-adjustable positioning***Attendant controlled Trendelenburg foot release (on left & right side)***Left side fold down table***Right side Pivot Table***Rear push handle**Lay Flat***Hide-away foot rest***Independent leg rest***LiquiCell for added comfort***1.800.237.3377***5516 SW 1st Lane***Ocala; Florida 334474***www.wincomfg.com***". The Care Cliner is a multi position recliner with a steel frame used primarily in clinics; hospitals and treatment centers for recovery; dialysis; oncology; examination; infusion. Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala; FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"Note: Model 6988 is a version of the 6980 without the footplate. We have a single; but high volume customer; that requested the chair without the footplate. We put the model in the system so as to call out a specific Bill of Materials to leave off the footplate. Since it was a unique request; there is not a piece of literature for it. | 2 | Apr-25-2012 | Winco Mfg.; LLC | Winco Manufacturing LLC is recalling the Care Cliner; XL Care Cliner; Drop Arm Care Cliner; Elite Care Cliner; and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this
Z-1458-2012 | Care Cliner - CARE CLINER 6550 (STD)/6570 (X-LG) with Drop Arm products are labeled in part: "***CARE CLINER 6550/6570 (w/ NYLON CASTERS)***Winco ISO 9001-2008 Certified***All of our best features; plus the added versatility of drop arms***Leave it to Winco to combine the best clinical chair features in the industry with the convenience of dual drop-arms! Our drop-arm care cliners make assisting your patients easier. Our side-panels lift; making clean-ups quick and easy.***KEY FEATURES***Standard head-rest cover***Padded drop-down armrests***Dual fold down side-tables***Easy-to-clean side panels***Trendelenburg Positioning***1.800.237.3377***5516 SW 1st Lane***Ocala; Florida 334474***www.wincomfg.com***". Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala; FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"The Care Cliner is a multi position recliner with a steel frame used primarily in clinics; hospitals and treatment centers for recovery; dialysis; oncology; examination; infusion. | 2 | Apr-25-2012 | Winco Mfg.; LLC | Winco Manufacturing LLC is recalling the Care Cliner; XL Care Cliner; Drop Arm Care Cliner; Elite Care Cliner; and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this
Z-1460-2012 | Care Cliner - ELITE CARE CLINER 6940(STD)/6950(X-LG) w/ Swing Arm products are labeled in part: "****ELITE CARE CLINER 6940/6950***Winco***ISO 9001-2008 Certified***Hassle-free design to optimize your patients comfort***This Elite Care Cliner features dual swing-arms; allowing patients with limited mobility a modified entry. Also; LiquiCell has been added into the seat for extra comfort. This ultra-thin; liquid-filled interface aids in improved blood flow; and helps prevent skin breakdown.***HEAT MASSAGE(Optional)***KEY FEATURES***Dual side swing-arms for modified entry and ease of cleaning***Dual fold down tables***LiquiCell for added comfort***5" nylon total-lock casters***Built-in headrest***Headrest cover***1.800.237.3377***5516 SW 1st Lane***Ocala; Florida 334474***www.wincomfg.com***".Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala; FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"The Care Cliner is a multi position recliner with a steel frame used primarily in clinics; hospitals and treatment centers for recovery; dialysis; oncology; examination; infusion. | 2 | Apr-25-2012 | Winco Mfg.; LLC | Winco Manufacturing LLC is recalling the Care Cliner; XL Care Cliner; Drop Arm Care Cliner; Elite Care Cliner; and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this
Z-1459-2012 | Care Cliner - ELITE CARE CLINER 6900 (STD)/6910(X-LG) with Swing-Arm products are labeled in part: "***ELITE CARE CLINER 6900/6910***Winco***ISO 9001-2008 Certified***Design; comfort; durability; and value that exceed your expectations.***Attention to detail and concern for comfort and quality; standards that Winco recliners are known for; places the Elite Care Cliner among the best-selling chairs on the market today. LiquiCell; an ultra-thin liquid filled interface that aids in reducing skin pressure; is standard for the Elite Care Cliner.***HEAT MASSAGE (Optional)***KEY FEATURES***Dual fold down tables***LiquiCell for added comfort***5¿¿¿ nylon total-lock casters***Built-in headrest***Headrest cover***1.800.237.3377***5516 SW 1st Lane***Ocala; Florida 334474***www.wincomfg.com***".Sticker with serial number is labeled in part: "***Manufactured By***Winco***Ocala; FL 34474***www.wincomfg.com*** Model No. ***Serial No.***"The Care Cliner is a multi position recliner with a steel frame used primarily in clinics; hospitals and treatment centers for recovery; dialysis; oncology; examination; infusion. | 2 | Apr-25-2012 | Winco Mfg.; LLC | Winco Manufacturing LLC is recalling the Care Cliner; XL Care Cliner; Drop Arm Care Cliner; Elite Care Cliner; and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this
Z-1386-2012 | ACCU-CHEK Aviva Nano Blood Glucose Meter - ACCU-CHEK Aviva Nano Blood Glucose Meter No product labeling; beyond the labels applied to the meter; was provided with the meter as these meters were never intended to be distributed. A representative copy of product labeling that would have been included in a finished device kit is attached.The Accu-Chek Aviva Nano Blood Glucose meter is intended to be used with Accu-Chek Aviva test strips for the quantitative measurement of glucose (sugar) in fresh capillary whole blood samples drawn from the fingertips or palm. The system is intended to be used by a single person and should not be shared. The system is intended for self-testing outside the body (in vitro diagnostic use) by people with diabetes at home as an aid to monitor the effectiveness of diabetes control. The system should not be used for the diagnosis of or screening of diabetes or for neonatal use. Alternative site testing should be done only during steady-state times (when glucose is not changing rapidly). | 2 | Apr-09-2012 | Roche Diagnostics Operations; Inc. | Accu-Chek Aviva Nano Blood Glucose meters that were manufactured for the purpose of validating a manufacturing process for product to be distributed outside the US were distributed in the US through eBay by an unauthorized third party. These meters were not intended for distribution; were not verified to meet final product specifications and were not distributed to customers with the necessary pr
Z-1148-2012 | AdvanSync; an orthodontic expansion screw - AdvanSync is an orthodontic device intended for the orthodontic treatment of Class II malocclusions. The AdvanSync part numbers and lot numbers affected by this recall are as follows:Part Number: 600-4000 (Kits-include 10 Right 600-4000 and 10 Left Assemblies). Lot Numbers: 070990588; 070976869; 060971515; 060971514; 04094278DPart Number: 600-4040 (5-Pack; Left Assembly). Lot Numbers: 080901333; 080902356; 070981416; 060971299; 060976084Part Number: 600-4041 (5-Pack; Right Assembly). Lot Numbers: 080902357; 080901335; 070981417; 060971304; 060976079 | 2 | Mar-05-2012 | Ormco Corporation | Manufacturing Error: There is a potential for a braze joint failure in the upper telescopic rod assembly. Although an incident could occur in which the eyelet loses connection from the rod; there is no risk of swallowing or aspirating the appliance because it will still be attached to the patient's crown.
Z-1147-2012 | Titanium Surfix Alpha Screws - Titanium SURFIX ALPHA SCREWS; a component of the Newdeal Hallu Lock Plate System. Surfix Alpha Variable angle locking system exists in different diameters and lengths. The product can be sold as a replenishment part individually packaged sterile; or as one (1) unit which is kitted within an instrumentation tray at which point; the product is non-sterile until the tray in which it is kitted is sterilized a the designated facility. The product is indicated for use in fixation of fractures; osteotomies or arthrodesis of the first metatarso-phalangeal joint. | 2 | Mar-02-2012 | Integra LifeSciences Corp. | Manufacturing error. The part number which corresponds to the length dimension of : 14 mm. was incorrectly etched on to screws with the actual measurement of: 12 mm.
Z-1136-2012 | Orthos CM Bracket; Orthodontic Metal Bracket; - Brand name: Orthos CM Bracket; Orthodontic Metal Bracket;Part Numbers: 454-1210; 454-0210.The intended use of this device is a fixed attachment on a tooth which holds an archwire during orthodontic treatment. | 3 | Mar-01-2012 | Ormco Corporation | The recall was initiated because Ormco has confirmed that manufacturing error resulted in the Orthos CM Brackets being incorrectly manufactured with a torque of -9?. The packaging was labeled with +9? torque; however the brackets contained inside the packaging are actually Orthos CM Brackets with a torque of -9?.
Z-1057-2012 | Powerheart AED; CardioVive AED; GE Responder AED; Nihon Kohden CardioLife AED - Powerheart G3 Automated External Defibrillators (Non-Wearable) and include the following models: Powerheart 9300A; 9300E; 9300P; 9390A; 9390E; CardioVive 92532; 92533; Nihon-Kohden 9200G; 9231; and GE Responder 2019198; 2023440 AEDs.The Powerheart AED G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation; if the victim is breathing; the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs; the device will charge automatically and be able to deliver therapy. The device is intended to be used by personnel who have been trained it is operation. The user should be qualified by training in basic life support or other physician-authorized medical response. | 1 | Feb-27-2012 | Cardiac Science Corporation | The affected Automated External Defibrillators (AEDs) contains a circuit board manufactured with a component that may fail unexpectedly due to a supplier manufacturing defect. If the component were to fail during a rescue attempt; the AED may not deliver defibrillation therapy.
Z-0803-2012 | Salter Labs Bubble humidifier - Salter Labs Bubble humidifier; with 6 PSI safety valve; 350cc; Latex Free; Single Patient use; used with supplemental oxygen therapy. Model Numbers 7600-0-50; E7600-0-50; NP7600-0-50 and special order patient kits that contain 7600 humidifiers. Patient kits have brand names of Salter Labs; AdvaCare; Allcare Medical; Memorial Home Services; Therapy Support Inc; Mendo-Lake Home Respiratory Services; MedAssurance Inc.; Abundant Home Care.Product Usage: Humidifier - single patient use. used with supplemental oxygen therapy. | 1 | Jan-26-2012 | SALTER LABS | A manufacturing defect was discovered internally which indicates that in a small number of humidifier lids; the pop-off may not occur at the appropriate pressure and can result in a leak.
Z-0512-2012 | Urine Reagent Strips (URS) 10 Parameters - Urine Reagent Strips (URS) 10 ParametersURS-l0-1OOW510(K) number: K970250Devise Listing number: D072377Each device consists of 100 strips of Urine Reagent Strips 10 parameterspacked with a desiccant pouch inside a gray bottle with Teco DiagnosticsLabel on the bottle and the bottle placed inside a Teco Diagnostics Box.For the semi-quantitative and qualitative detection of Glucose; Bilirubin; Ketone; Specific Gravity; Blood; pH; Protein; Urobilinogen; Nitrite; Leukocytes in urine. | 2 | Jan-11-2012 | Teco Diagnostics | This recall was initiated in response to findings during a FDA inspection of the firm. The firm's recall has been initiated due to improper manufacture of sub-component nitrite paper Lot 35134; Expiration Date: 01/2014. Discrepancies were found on the information recorded on formulation sheet and manufacturing log sheets (Refer to next paragraph). Sampling size submitted for QC testing was insuff
Z-0511-2012 | Urine Reagent Strips (URS) 10 Parameters - Urine Reagent Strips (URS) 10 ParametersURS-l0-1OOW510(K) number: K970250Devise Listing number: D072377Each device consists of 100 strips of Urine Reagent Strips 10 parameterspacked with a desiccant pouch inside a gray bottle with Teco DiagnosticsLabel on the bottle and the bottle placed inside a Teco Diagnostics Box.For the semi-quantitative and qualitative detection of Glucose; Bilirubin; Ketone; Specific Gravity; Blood; pH; Protein; Urobilinogen; Nitrite; Leukocytes in urine. | 2 | Jan-11-2012 | Teco Diagnostics | This recall was initiated in response to findings during a FDA inspection of the firm. The firm's recall has been initiated due to improper manufacture of sub-component nitrite paper Lot 35134; Expiration Date: 01/2014. Discrepancies were found on the information recorded on formulation sheet and manufacturing log sheets (Refer to next paragraph). Sampling size submitted for QC testing was insuff
Z-0463-2012 | Smith& Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE Resector - Smith & Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE ResectorCatalog Number: 7205962. | 2 | Jan-11-2012 | Smith & Nephew; Inc. Endoscopy Division | Manufacturing issue: The distal tip of the outer sheath (return electrode) may detach and become loose in the joint during use.
Z-0296-2012 | Respironics Trilogy 100; 200 and 202 Ventilators - Trilogy 100; 200 and 202 VentilatorsThe device is intended to be used in home; institution/hospital; and portable applications such as wheelchairs and gurneys; and may be used for both invasive and non-invasive ventilation. | 1 | Jan-11-2012 | Respironics; Inc. | Trilogy 100; 200 and 202 units are being recalled due to low impeller press-force used during the manufacturing process. If the press force were low enough; theoretically the impeller could move up the shaft and come into contact with the inside of the blower housing/impeller enclosure. This may impact therapy.
Z-0295-2012 | GammaMed Flexible Applicator Probe - GammaMed Flexible Applicator Probe; Part GM11002560 - 3.2 mm diameter braced PVDF. Varian Medical Systems; Palo Alto; CA; Mfg. by Varian Medical Systems; Haan GmbH. Designed for brachytherapy intracavity treatments in areas such as cancer of the vagina; vaginal stump or rectum. | 2 | Nov-30-2011 | Varian Medical Systems; Inc. | GammaMed Flexible Applicator Probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave.
Z-0271-2012 | TUM-E-VAC - TUM-E-VAC REF 2075; Rx ONLY; NON STERILE --- Ethox International; Buffalo; NY 14204; USA --- Intended use: gastric lavage | 3 | Nov-22-2011 | Ethox International; Inc. | Ethox Tum-E-Vac Gastric Lavage Kit #2075; Lot Number 031123255 was labeled with expiration date symbol and date of 2010/10 on the unit product label. The carton was labeled with a manufacturing date symbol and date of 2010/10. The Ethox Tum-E-Vac Gastric Lavage Kit #2075 does not have an expiration date.
Z-0270-2012 | Uplift Commode Assist - "***UPLIFT TECHNOLOGIES INC.***Bringing quality to life.***Model No.***CA200***Description***Lifting commode chair***Weight Range***80-300+ lb-136+ kg***"Labeling on product: "***1C0101 and K0061014***".A lifting commode chair that is self-powered (non-electric) and provides lifting assistance for those who have reduced mobility due to arthritis; post-operative conditions; degenerative joint disease; Parkinson's disease or other neuromuscular disorders. | 2 | Nov-22-2011 | Uplift Technologies Inc. | On June 26; 2011 Uplift Technologies Inc. Nova Scotia; Canada initiated a recall of the Uplift Commode Assist Model No. CA200; Lot No. 1C0101. The plastic seat does not meet manufacturing specifications. The seat may not fit the unit properly and may interfere with the unit's normal locking mechanism.
Z-0200-2012 | Spirit MB Brackets - Spirit MB Brackets; part #494-1210.The brand name of the device is Spirit MB Brackets; an orthodontic bracket. Spirit MB Brackets have been assigned the product code 76 DYW (Plastic Orthodontic Bracket) by the FDA and is classified as a Class II Medical Device. Spirit MB Brackets are to be fixed attachment on a tooth which holds an archwire during orthodontic treatment. | 3 | Nov-16-2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error.
Z-0201-2012 | Spirit MB Brackets - Spirit MB Brackets; part #494-1211.The brand name of the device is Spirit MB Brackets; an orthodontic bracket. Spirit MB Brackets have been assigned the product code 76 DYW (Plastic Orthodontic Bracket) by the FDA and is classified as a Class II Medical Device. Spirit MB Brackets are to be fixed attachment on a tooth which holds an archwire during orthodontic treatment. | 3 | Nov-16-2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error.
Z-0171-2012 | OSSEOTITE Certain Implant - ***REF IOSS511***OSSEOTITE Certain Implant 5 x 11.5mm***LOT 910695***Sterile using Radiation***Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park; Ed; 1; Planta 1 Pl. de la Pau s/n 08940 Cornella de Llobregat (Barcelona) Spain***Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment. The Thread form titanium screws (cylindrical dental implants) designed to be placed into mandibular or maxillary jaws in order to affix a dental prosthesis (edentulous or partially edentulous segments) that replaces 1 or more missing teeth. | 2 | Nov-08-2011 | Biomet 3i; LLC | The product may contain a manufacturing defect; that if present; may prevent the driver; cover screw or abutment from fully engaging the implant.
Z-3279-2011 | Zimmer Trabecular Metal Modular Acetabular System - The Trabecular Metal Modular Acetabular System is a modular acetabular cup system consisting of a shell with a preassembled locking ring; a liner; and optional screws. The Shells are available in a variety of sizes and are designed to be used with Trilogy Acetabular System liners. The shell implants are used to replace the hip socket during total hip arthroplasty.The device is indicated for primary or revision surgery for rehabilitating hips damaged as a result on noninflammatory degenerative joint disease (NIDJD) or its composite diagnoses of osteoarthritis; avascular necrosis;protrusio acetabuli; traumatic arthritis; slipped capital epiphysis; fused hip; fracture of the pelvis and diastrophic variant. The device is intended for either cemented or noncemented use. | 2 | Sep-27-2011 | Zimmer Inc. | Lot # 61132698 and Lot 001122467 (added 9/ 18/2012) may have a manufacturing non conformance that could prevent the devices from functioning as intended. Specifically; devices in the affected lot may have a locking ring groove depth that prevents the corresponding liner from snapping into the shell. The groove provides the locking ring room to displace when inserting a liner into the shell. Th
Z-3177-2011 | Monomid Non-Absorbable Suture - Monomid Non-Absorbable Suture are blue; monofilament synthetic absorbable suture; with different sizes and different needle sizes.The products are labeled in parts.These sutures are intended for use in general soft tissue approximation and/or ligation including use in cardiovascular; ophthalmic; and neurological surgery. | 2 | Sep-09-2011 | C P Medical Inc. | Unapproved dye was used in manufacturing the non-absorbable blue nylon suture material.
Z-3013-2011 | GammaMed Flexible Applicator - GammaMed Flexible Applicator probe; 3.2mm diameter; braced PVDF.Varian Medical Systems; Palo Alto; CA.Mfg by Varian Medical Systems; Haan GmbH.Flexible Applicator Robes are designed for brachytherapy intracavitary treatments in areas such as vagina; vagina stump or rectum. | 2 | Aug-12-2011 | Varian Medical Systems; Inc. Oncology Systems | Certain GammaMed Flexible applicator probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave.
Z-2946-2011 | Vascular Solutions; Pronto¿ V4; Extraction Catheter - Vascular Solutions; Pronto¿ V4; Extraction Catheter; Sterilized using ethylene oxide; Rx Only. The following sizes/model numbers are affected: 5.5F( 4005); 6F(4006; 7F(4007; 8F(4008) are involved in the recall.The Pronto catheter is indicated for the removal of fresh; soft emboli and thrombi from vessels in the coronary and peripheral vasculature. | 2 | Aug-04-2011 | Vascular Solutions; Inc. | Vascular Solutions ; Inc recently discovered that there was a manufacturing defect with Pronto V4 extraction catheters (Models 4005; 4006; 4007; 4008). Use of this defective product could lead to a tip separation. This maycause injury and/or procedural delay; resulting in a deteriorating hemodynamic condition of the patientwhich may require medical intervention.
Z-2838-2011 | Incise Zirconia Coping (CBZR0102) - "***3i Incise Custom-Made Label. QP157 Approved 07/09/10*** 3i inc se Custom-made Dental Framework***Biomet 3i 4555 Riverside Drive Palm Beach Gardens; FL 33410 800-342-5454 www.biomet3i.com Please consult instructions for use and recommendation***For Dental Lab use ONLY Non-sterile/For single Use***"Indicated for use in restorative procedure to affix the final prosthesis to the abutment. Coping and frameworks provide the base structure for final porcelain application as the final step of creating a restoration. The devices are intended to be used to create final restoration for dental reconstruction with either a prepared tooth or dental implant abutment. | 3 | Jul-16-2011 | Biomet 3i; LLC | On December 1; 2010 Biomet 3i recalled their product Incise Zirconia Coping (CBZR0102). The Coping contains a manufacturing defect. Coping is oversized. All affected direct accounts were notified by letter on December 14; 2010.
Z-2618-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit; Part number 65418-01; Terumo Latin America Seminsa Pediatric Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2615-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit Part number 70667 Eisenhower Army Med Center Heart Lung Pack; Sterile; Terumo Cardiovascular Systems Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2613-2011 | Pressure Monitoring Kit - Pressure Monitoring Kit; Catalog 16066100; Terumo Cardiovascular Systems Ann Arbor MI. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2620-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit; Part number 65870; Terumo Latin America Pediatric Pack; Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2621-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit; Part number 65871; Terumo Latin America Lactante Pack; Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2616-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit Part number 70367-07 Mission St. Joe's - Asheville X-Coated Low Prime Pack; Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2619-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit; Part number 65672-01; Terumo Latin America Seminsa lactante; Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2614-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit Part number 70667-01 Eisenhower Army Med Center Heart Lung Pack; Sterile; Terumo Cardiovascular Systems Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2617-2011 | Pressure Monitoring Kit - Cardiovascular Procedure Kit Part number 72427; Oklahoma Heart Hospital X-Coated System 1 Pack; Sterile; Terumo Cardiovascular Systems; Ashland; MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. | 2 | Jun-22-2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process.
Z-2383-2011 | PTFE Coated 3 cm Flexible Tip Guidewire - Gyrus ACMI; REF 25 BX; PTFE Coated 3 cm Flexible Tip Guidewire; .035" x 150 cm; Sterile EO; Rx Only; Gyrus ACMI; Inc. 136 Turnpike Road;Southborough; MA 01772-2104 USAUreteral guidewires are used to establish access in the urinary system for the passage of urological devices. An introducer is first used to establish entry to various portions of the urinary tract (urethra; bladder; ureter); then a guidewire is threaded through the introducer. Following placement of the guidewire; various urological devices can be introduced into the targeted portion of the urinary tract for various therapeutic and diagnostic procedures. | 2 | May-26-2011 | Gyrus Acmi; Incorporated | Gyrus ACMI; Inc. wants to bring a recall issue to their customers attention concerning one lot of their 25BX-PTFE Coated 3 cm Flexible Tip Guidewire. As a result of a manufacturing error in the production this lot of guidewires; some units of product were loaded in the dispenser backwards. In such cases; the stiff end of the guidewire would be dispensed from the product first and; if undetected;
Z-2054-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 32 MM O.D. 55 MM O.D.; REF 71338950; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06 | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-2056-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 36 MM O.D. 54 MM O.D.; REF 71338954; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06 | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-2053-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 32 MM I.D. 48 MM O.D.; REF 71338948; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06Usage: Primary total hip arthroplasty due to non-inflammatory arthritis (degenerative joint disease) such as osteoarthritis; avascular necrosis or traumatic arthritis | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-2055-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 36 MM O.D. 52 MM O.D.; REF 71338952; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06 | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-2057-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 36 MM O.D. 56 MM O.D.; REF 71338956; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06 | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-2058-2011 | R3 Forte Ceramic Liners - R3 TM CERAMIC ACETABULAR LINER; BIOLOX¿ FORTE; 36 MM O.D. 62 MM O.D.; REF 71338962; STERILE R; QTY (1); Smith & Nephew; Inc.; Memphis; TN 38116; 2017/06 | 2 | Apr-22-2011 | Smith & Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners; the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners.
Z-1966-2011 | Container; IV - ExactaMix EVA Container; 2000mL; REF: 740; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1962-2011 | Container; IV - ExactaMix EVA Container; Calibration Bag; REF: 735; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1967-2011 | Container; IV - ExactaMix EVA Container; 3000mL; REF: 741; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1964-2011 | Container; IV - ExactaMix EVA Container; 500mL; REF: 738; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1965-2011 | Container; IV - ExactaMix EVA Container; 1000mL; REF: 739; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1963-2011 | Container; IV - ExactaMix EVA Container; 250mL; REF: 737; Sterile R; Baxa Corporation; Englewood; CO 80112.A plastic container used to hold large-volume sterile solutions to be administered to a patient through an intravascular administration set | 2 | Apr-12-2011 | Baxa Corporation | TPN bags may leak during filling operation; due to bag manufacturing error.
Z-1690-2011 | DePuy CEMENT RESTRICTOR; Size 5 - DePuy; a Johnson & Johnson company; CEMENT RESTRICTOR; Size 5; Sterile; DePuy Orthopaedics; 700 Orthopaedic Drive; Warsaw; IN; 46582; 800-366-8143/This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis | 2 | Mar-17-2011 | Depuy Orthopaedics; Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised; the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field; including the implant. Non-sterile product and/or contaminat
Z-1687-2011 | DePuy CEMENT RESTRICTOR; Size 4 - DePuy; a Johnson & Johnson company; CEMENT RESTRICTOR; Size 4; Sterile; DePuy Orthopaedics; 700 Orthopaedic Drive; Warsaw; IN; 46582; 800-366-8143/This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis | 2 | Mar-17-2011 | Depuy Orthopaedics; Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised; the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field; including the implant. Non-sterile product and/or contaminat
Z-1688-2011 | DePuy CEMENT RESTRICTOR; Size 5 - DePuy; a Johnson & Johnson company; CEMENT RESTRICTOR; Size 5; Sterile; DePuy Orthopaedics; 700 Orthopaedic Drive; Warsaw; IN; 46582; 800-366-8143/This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis. | 2 | Mar-17-2011 | Depuy Orthopaedics; Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised; the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field; including the implant. Non-sterile product and/or contaminat
Z-1689-2011 | DePuy CEMENT RESTRICTOR; Size 3 - DePuy; a Johnson & Johnson company; CEMENT RESTRICTOR; Size 3; Sterile; DePuy Orthopaedics; 700 Orthopaedic Drive; Warsaw; IN; 46582; 800-366-8143/This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis. | 2 | Mar-17-2011 | Depuy Orthopaedics; Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised; the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field; including the implant. Non-sterile product and/or contaminat
Z-1686-2011 | DePuy CEMENT RESTRICTOR ; Size 1 - DePuy; a Johnson & Johnson company; CEMENT RESTRICTOR; Size 1; Sterile; DePuy Orthopaedics; 700 Orthopaedic Drive; Warsaw; IN; 46582; 800-366-8143/This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis. | 2 | Mar-17-2011 | Depuy Orthopaedics; Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised; the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field; including the implant. Non-sterile product and/or contaminat
Z-0906-2011 | InfraVision IR Illuminator Console - InfraVision IR Illuminator Console; model number 220-180-521; manufactured by Stryker Endoscopy San Jose; for use with the Stryker InfraVision system.Intended use:Medical device for human use; during surgical procedures to transilluminate anatomical structures during surgical procedures with infrared light. It can be used in either open surgical interventions or during laparoscopic procedures. infrared light which is not visible to the human eye is detected by all of Stryker's infrared sensitive cameras. The InfraVision console outputs the infrared light via disposable fibers hooked to the console that are contained in either the E-kit or the U-kit. | 2 | Jan-18-2011 | Stryker Endoscopy | Product may not adequately visualize IR light emitted from U- and E- kits due to a manufacturing error.
Z-0824-2011 | Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long - Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long; Catalog 00-2490-003-50; Zimmer; Warsaw; IN.A reusable surgical instrument used for the insertion of a Lag Screw used in the Zimmer Natural Nail system. | 2 | Dec-28-2010 | Zimmer Inc. | The firm has determined that two lots of this instrument were inappropriately manufactured by the supplier. The manufacturing process created a step in the lumen which impedes the Lag Screw Retaining Shaft from passing through the lumen without the use of excessive force.
Z-0831-2011 | Manual Wheelchairs with Transit System Option - Manual Wheelchairs with Transit System Option;Model: Quickie LXI (EILXI);Product is manufactured and distributed by Sunrise Medical; Inc. Fresno; CA Intended to provide mobility to persons restricted to a sitting position and are unable to transfer from their wheelchairs when traveling in motor vehicles; the wheelchair must serve as a vehicle seat. | 2 | Dec-28-2010 | Sunrise Medical Inc | The stabilizer bar was inadvertently omitted during the manufacturing process.
Z-0545-2011 | Allura XPER - Allura XPER X-ray system.The intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0549-2011 | ALLURA XPER FD20/10 - ALLURA XPER FD20/10The intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0544-2011 | MultiDiagnost Eleva with FD - MultiDiagnost Eleva with FD.The Philips MultiDiagnost Eleva intended use is for the following applications: As a diagnostic-imaging device multi-functional/ universal system; general R/F; Fluoroscopy; Radiography and Angiography can be performed along with pediatric examinations and some more specialized interventional applications. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0548-2011 | ALLURA XPER FD20 - ALLURA XPER FD20The intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0547-2011 | ALLURA XPER FD10/10 - ALLURA XPER FD10/10The intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0546-2011 | Allura XPER FD10 F/C - Allura XPER FD10 F/CThe intended uses of the Allura XPER X-ray systems involved with this report are angiographic X-ray systems with a solid state x-ray imaging device for cardiovascular and vascular diagnostic and interventional procedures. | 2 | Dec-07-2010 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result; the complete system may shut down and not re-start.
Z-0482-2011 | Neuroform 3 Microdelivery Stent System - Neuroform 3 Microdelivery Stent System; Distributed and manufactured by: Boston Scientific Corporation; Fremont; CA | 2 | Nov-26-2010 | Boston Scientific Corporation | The product may contain manufacturing aids (mandrels) which should not be present in the final product. If the mandrel remains in place; the catheter cannot be loaded over the required guidewire; so cannot be used.
Z-0308-2011 | Panta Arthrodesis Nail - Panta Support Device Panta¿ Arthrodesis Nail Instrumentation Tray One unit per package.The PANTA(R) Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. | 2 | Nov-02-2010 | Integra LifeSciences Corporation | Inability to properly align the calcaneal screw placement on a consistent basis; due to a manufacturing defect of of the targeting frame (Panta Support Device).
Z-2478-2010 | Fabius GS - Fabius GS Anesthesia Machine; OR/Anesthesia; Draeger Medical; Inc.; 3135 Quarry Road; Telford; PA 18969inhalation anesthesia machine for use in operating; induction and recovery rooms | 2 | Sep-23-2010 | Draeger Medical Systems; Inc. | potential failures to Electri-Cord Manufacturing AC power cord
Z-2479-2010 | Fabius Tiro - Fabius Tiro Anesthesia Machine; OR/Anesthesia; Draeger Medical; Inc.; 3135 Quarry Road; Telford; PA 18969Compact anesthesia system that offers the full spectrum of anesthesia ventilation to military anesthesiologists/ CRNAs in the field of operations. Can be used in a variety of mobile and stationary military applications where general anesthesia is required. | 2 | Sep-23-2010 | Draeger Medical Systems; Inc. | potential failures to Electri-Cord Manufacturing AC power cord
Z-2482-2010 | ARCHITECT¿ i1000SR" Processing Module - ARCHITECT¿ i1000SR" Processing Module System For In Vitro Diagnostic Use; Abbott Laboratories Abbott Park; ILProduct List Number; Part # or Model: 01L86-01In Vitro Diagnostic use only. The Abbott ARCHITECT System is designed to perform automated: Chemistry tests; utilizing photometry and potentiometric technology; and Immunoassay tests; utilizing CMIA (chemiluminescent microparticle assay) detection technology. | 2 | Sep-23-2010 | Abbott Laboratories; Inc | During manufacturing; two components of the ARCHITECT i1000SR were reversed which could potentially cause falsely elevated patient results due to a sample-reagent and reagent-sample carryover.
Z-2444-2010 | St. Jude Medical; Model 3085; Dual-Chamber External Pulse Generator - St. Jude Medical; Model 3085; Dual-Chamber External Pulse GeneratorThe PACE 203H / Model 3085 external pulse generator / temporary cardiac pacemaker is designed to be used with cardiac pacing lead systems for temporary atrial or ventricular or A-V sequential pacing. The PACE 203H / Model 3085 has applications where such pacing modes are indicated for therapeutic; prophylactic; or diagnostic purposes. | 2 | Sep-21-2010 | Osypka Medical Inc | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e.; not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t
Z-2443-2010 | Oscor PACE 203H; Dual-Chamber External Pulse Generator - Oscor PACE 203H; Dual-Chamber External Pulse GeneratorThe PACE 203H / Model 3085 external pulse generator / temporary cardiac pacemaker is designed to be used with cardiac pacing lead systems for temporary atrial or ventricular or A-V sequential pacing. The PACE 203H / Model 3085 has applications where such pacing modes are indicated for therapeutic; prophylactic; or diagnostic purposes. | 2 | Sep-21-2010 | Osypka Medical Inc | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e.; not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t
Z-2436-2010 | Pleatman Sac¿ Tissue Removal System - The Gyrus ACMI Pleatman Sac Tissue Removal System facilitates removal of specimens while minimizing contamination of the abdominal cavity or wound tract. The system contains one (1) Sac and one (1) obturator in a sterile; ready -to -use; single- use procedure tray. | 2 | Sep-19-2010 | Gyrus Acmi; Incorporated | Gyrus Medical; Inc. has issued a Urgent Medical Device Recall for the Gyrus ACMI product 004942-903 - Pleatman Sac? Tissue Removal System; sold as lot 186540FC with an expiration date of 06/2015. As a result of a manufacturing error in the production of the trays used in this lot; it is possible that a crack could form that would compromise the package integrity; and possibly the sterility of th
Z-2242-2010 | CIVCO Surgi-Tip Transducer cover - CIVCO Surgi-Tip Transducer Cover; Ref. 610-834; Sterile (20.3 x 244cm) polyethylene telescopically-folded cover with attached (6cm) NeoGuard tip. The latex-free surgi tipped covers are utilized to cover ultrasound transducers and prevent cross-contamination from occurring with equipment and patients. | 2 | Aug-20-2010 | Civco Medical Instruments Inc | Polyethylene material did not have proper treatment applied during manufacturing. Untreated polyethylene material does not maintain a secure seal with the adhesive tape used to secure the tip to the cover.
Z-2212-2010 | DYNASTY ¿ A-CLASS ¿ POLY LINER - DYNASTY¿ A-CLASS¿ POLY LINER; REF DLXP-LD36; 1 EACH; Rx ONLY; STERILE EO; GROUP D; I.D. 36 mm; LIP 15o; LINER STD; Wright Medical Technology; 4677 Airline Road; Arlington; TN 38002. Product is used in total hip arthroplasty. | 3 | Aug-12-2010 | Wright Medical Technology Inc | The inner and outer product packaging was missing the following information: shelf life; translations; manufacturing date; CE marking and also listed the incorrect sterilization method.
Z-2205-2010 | Pulmonetic Systems - Pulmonetic Systems P/N 19189-001 Patient Circuit SPU W/O PEEP 15 FT 5/Pkg.These circuits are intended for use with the MR Conditional LTV 1200 System with the ventilator utilizing an external exhalation valve; as single patient use (SPU) only patient circuits for adult and pediatric patients. The LTV 1200 ventilator is intended to provide continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation. The ventilator is suitable for use in institutional; home and transport settings; and is applicable for adult and pediatric patients weighing at least 5 kg (11 lbs.); who require the following types of ventilatory support: -Positive Pressure Ventilation; delivered invasively (via ET tube) or non-invasively (via mask). -Assist/Control; SIMV; CPAP or NPPV modes of ventilation. -Breath types including Volume; Pressure Control and Pressure Support. | 2 | Aug-10-2010 | CareFusion 203; Inc. | CareFusion is recalling the P/N 19189-001 Patient Circuit (SPU W/O PEEP 15ft) because some of the circuit sense lines may have been reversed during the manufacturing process. If the sense lines have been reversed; immediately upon connection of the patient circuit to the ventilator and turning the ventilator on; the ventilator will begin to autocycle (with increased pressure delivered) with both
Z-1767-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 55 cm Straight Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 55 cm Straight Basic Kit; Catalog Number/REF 10302809; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1762-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 24 cm Straight Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 24 cm Straight Basic Kit; Catalog Number/REF 10302803; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1771-2010 | DURAMAX Hemodialysis Catheters; Stack Tip 36 cm Pre-Curve Basic Kit - DURAMAX (TM) Hemodialysis Catheters; Stack Tip 36 cm Pre-Curve Basic Kit; Catalog Number/REF 10302815; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1764-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 32 cm Straight Basic Kit - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 32 cm Straight Basic Kit; Catalog Number/REF 10302805; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1765-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 36 cm Straight Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 36 cm Straight Basic Kit; Catalog Number/REF 10302806; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1768-2010 | DURAMAX Hemodialysis Catheters; Stack Tip 24 cm Pre-Curve Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stack Tip 24 cm Pre-Curve Basic Kit; Catalog Number/REF 10302812; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1770-2010 | DURAMAX Hemodialysis Catheters; Stack Tip 32 cm Pre-Curve Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stack Tip 32 cm Pre-Curve Basic Kit; Catalog Number/REF 10302814; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1761-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 22 cm Straight Basic Kit - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 22 cm Straight Basic Kit; Catalog Number/REF 10302802; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1766-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 40 cm Straight Basic Kit - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 40 cm Straight Basic Kit; Catalog Number/REF 10302807; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1763-2010 | DURAMAX Hemodialysis Catheters; Stacked Tip 28 cm Straight Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stacked Tip 28 cm Straight Basic Kit; Catalog Number/REF 10302804; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1769-2010 | DURAMAX Hemodialysis Catheters; Stack Tip 28 cm Pre-Curve Basic Kit; - DURAMAX (TM) Hemodialysis Catheters; Stack Tip 28 cm Pre-Curve Basic Kit; Catalog Number/REF 10302813; 510 (k) 080400; Sterile --- Manufactured in USA by AngioDynamics; Inc. 603 Queensbury Avenue; Queensbury; NY 12804 | 2 | Jun-04-2010 | Angiodynamics; Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test.
Z-1681-2010 | Lyric Hearing Aid - InSound Medical; Inc; Lyric in-canal hearing aid. | 2 | May-26-2010 | InSound Medical; Inc. | Manufacturing error could result in electrolyte leakage from the product's battery.
Z-1611-2010 | 4.5MM Bone Tap; Modular - 4.5MM Bone Tap; Modular; Manufactured by: Blackstone Medical; Inc. 1720 Bray Central Drive; McKinney; TX 75069.Intended to created a screw thread in a hole drilled into bone to facilitate placement of a bone screw. | 2 | May-19-2010 | Blackstone Medical; Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread.
Z-1614-2010 | 7.5MM Bone Tap; Modular - 7.5MM Bone Tap; Modular; Manufactured by: Blackstone Medical; Inc. 1720 Bray Central Drive; McKinney; TX 75069.Intended to created a screw thread in a hole drilled into bone to facilitate placement of a bone screw. | 2 | May-19-2010 | Blackstone Medical; Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread.
Z-1615-2010 | 8.5MM Bone Tap; Modular - 8.5MM Bone Tap; Modular; Manufactured by: Blackstone Medical; Inc. 1720 Bray Central Drive; McKinney; TX 75069.Intended to created a screw thread in a hole drilled into bone to facilitate placement of a bone screw. | 2 | May-19-2010 | Blackstone Medical; Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread.
Z-1613-2010 | - 6.5MM Bone Tap; Modular; Manufactured by: Blackstone Medical; Inc. 1720 Bray Central Drive; McKinney; TX 75069.Intended to created a screw thread in a hole drilled into bone to facilitate placement of a bone screw. | 2 | May-19-2010 | Blackstone Medical; Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread.
Z-1612-2010 | 5.5MM Bone Tap; Modular - 5.5MM Bone Tap; Modular; Manufactured by: Blackstone Medical; Inc. 1720 Bray Central Drive; McKinney; TX 75069.Intended to created a screw thread in a hole drilled into bone to facilitate placement of a bone screw. | 2 | May-19-2010 | Blackstone Medical; Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread.
Z-1599-2010 | Eon Mini" Implantable Pulse Generator (IPG) - Eon Mini Implantable Pulse Generator (IPG). Advanced Neuromodulation Systems; Inc. Plano; TX.Indicated as an aid in the management of chronic; intractable pain of the trunk and/or limbs; including unilateral or bilateral pain. | 2 | May-13-2010 | Advanced Neuromodulation Systems; Inc | Due to a missed inspection step during manufacturing; a small number of devices may experience a gradual or sudden loss of charging capability.
Z-1559-2010 | VENTAK PRIZM Implantable Cardioverter Defibrillator - VENTAK PRIZM Automatic Implantable Cardioverter Defibrillator (AICD). DR; model 1851; DR HE; model 1853 and 1858; PRIZM VR HE; model 1857; PRIZM 2 VR model 1860 and PRIZM 2 DR model 1861. The ICDs are designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation (VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-energy shocks using either a biphasic or monophasic waveform. Bradycardia pacing; including adaptive-rate features is available to detect and treat bradyarrhythmias and to support the cardiac rhythm after defibrillation therapy. The devices denoted with DR offer dual-chamber bradycardia features (atrial and/or ventricular pacing and sensing); and the devices denoted with VR offer single-chamber bradycardia features (ventricular pacing and sensing). Sterilized with gaseous ethylene oxide. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1555-2010 | CONFIENT Implantable Cardioverter Defibrillator - Boston Scientific CONFIENT; model E030; Implantable Cardioverter Defibrillator (ICD). The CONFIENT Implantable Cardioverter Defibrillator (ICDs) are designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation(VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-energy shocks using either a biphasic or monophasic waveform. Sterilized using ethylene oxide. Boston Scientific; 4100 Hamline Avenue; St. Paul; MN 55112-5798. USA. | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1553-2010 | TELIGEN Implantable Cardiac Defibrillator - Boston Scientific TELIGEN; Implantable Cardiac Defibrillator; Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific; 4100 Hamline Avenue North; St. Paul; MN 66112-5798; USA. The product is indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. COGNIS devices also provide bi-ventricular pacing. | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1554-2010 | COGNIS CRT-D; Cardiac Resynchronization Therapy Defibrillator; - Boston Scientific COGNIS CRT-D; Cardiac Resynchronization Therapy Defibrillator; Models N118 and N119. Sterilized using ethylene oxide. Boston Scientific; 4100 Hamline Avenue North; St. Paul; MN 66112-5798; USA. The product is indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. COGNIS devices also provide bi-ventricular pacing. | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1558-2010 | CONTAK RENEWAL Cardiac Resynchronization Therapy Defibrillator - CONTAK RENEWAL 3 (Models H170); Guidant CONTAK RENEWAL 3 HE (Models H177; H179). Cardiac Resynchronization Therapy Defibrillator (CRT-D). These CRT-Ds provide ventricular tachyarrhythmia and cardiac resynchronization therapies. Ventricular tachyarrhythmia therapy is for the treatment of ventricular tachycardia (VT) and ventricular fibrillation (VF); rhythms that are associated with sudden cardiac death (SCD). Cardiac resynchronization therapy is for the treatment of heart failure (HF) and uses biventricular electrical stimulation to synchronize ventricular contractions. Cardioversion/defibrillation therapies include a range of low- and high-energy shocks using either a biphasic or monophasic waveform. Sterile EO. Boston Scientific Corporation; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1556-2010 | LIVIAN Cardiac Resynchronization Therapy Defibrillator - Boston Scientific LIVIAN Cardiac Resynchronization Therapy Defibrillator (CRT-D); models H220; H225; H227; H229. The LIVIAN family of cardiac resynchronization therapy defibrillators (CRT-Ds); provide ventricular tachyarrhythmia and cardiac resynchronization therapies. Ventricular tachyarrhythmia therapy is for the treatment of ventricular tachycardia (VT) and ventricular fibrillation (VF); rhythms that are associated with sudden cardiac death (SCD). Cardiac resynchronization therapy is for the treatment of heart failure(HF) and uses biventricular electrical stimulation to synchronize ventricular contractions. Sterilized using ethylene oxide. Boston Scientific; 4100 Hamline Avenue; St. Paul; MN 55112-5798 USA. | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1557-2010 | VITALITY Implantable Cardioverter Defibrillator - VITALITY Implantable Cardioverter Defibrillator. Atrial and Ventricular Therapies. AVT; model A155. VITALITY DS DR/VR (Models T125; T135); VITALITY 2 DR/VR (Models T165; T175) and VITALITY 2 EL DR/VR (Models T167; T177) Implantable Cardioverter Defibrillator. Guidant Corporation Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798; USA. | 2 | May-08-2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food & Drug Administration.
Z-1221-2010 | Point Scanning Confocal Microscope TCS SP5/TCS SP5 X - Shutter HX SP5; a component of the Leica Point Scanning Confocal Microscope Models TCS SP5 and TCS SP5 X.Laboratory instrument. | 2 | May-06-2010 | Leica Microsystems; Inc. | A manufacturing defect was identified that could potentially result in the laser shutter not closing as intended which could result in the shutter not completely closing disrupting the laser beam safely.
Z-1409-2010 | Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 - Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 | 2 | Apr-16-2010 | Boston Scientific Neuromodulation Corporation | Toshiba has determined that certain Sony battery packs installed in Toshiba portable computers can overheat; posing fire hazard to customers. Based on Toshiba's investigation with Sony Corporation; they concluded that the battery cells in specific manufacturing lots could be affected.
Z-1301-2010 | Spanish 100-120V 515B Pulse Oximeter - Spanish 100-120V 515B Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1283-2010 | Tidal Wave Model 615 - Tidal Wave Model 615; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1306-2010 | 515B Pulse Oximeter - 515B Pulse Oximeter; Refurbished - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1260-2010 | DeviceLink Connectivity Hardware - DeviceLink Connectivity Hardware.Supports Philips specified manufacturers' devices including urimeters; ventilators; oximeters; patient monitors; infusion pumps and anesthesia machines. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1303-2010 | 100-120V 2001 Pulse Oximeter with Mars Nurse Call - 100-120V 2001 Pulse Oximeter with Mars Nurse Call - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1296-2010 | Japanese 100-120V Oxypleth - 520A Pulse Oximeter - Japanese 100-120V Oxypleth - 520A Pulse Oximeter- Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1266-2010 | 100-120V Capnogard 1265 - 100-120V Capnogard 1265; Respironics/Philips product.Intended to be used for monitoring end tidal CO2 and respiration rate in all critical monitoring environments. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1272-2010 | Spanish 100-120V TCOM-860 - Spanish 100-120V TCOM-860; Respironics/Philips product.Used for trending O2 and CO2 tension at the skin surface for neonates and adults as an adjunct to arterial O2 and CO2 measurements. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1278-2010 | Tidal Wave Sp Model 715 - Tidal Wave Sp Model 715; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1277-2010 | Japanese TidalWave SP 710 - Japanese TidalWave SP 710; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1258-2010 | HeartStart MRx Defibrillator/Monitor - HeartStart MRx Defibrillator/Monitor.For use for the termination of ventricular tachycardia and ventricular fibrillation. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1269-2010 | Spanish 100-120V Cosmo - 7100 Capnograp - Spanish 100-120V Cosmo - 7100 Capnograp; Respironics/Philips products.Intended to be used for monitoring end tidal CO2; respiration rate; oxygen saturation and pulse rate in critical monitoring environments. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1274-2010 | 8100 Cosmo+ Resp Profile Monitor - Resp Profile Monitor; 8100 Cosmo+; Respironics/Philips product.Used to provide spirometric; carbon dioxide and pulse oximetry monitoring. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1282-2010 | Spanish 100-120V NICO 7300 - Spanish 100-120V NICO 7300; Respironics/Philips product.Intended use is cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation during general anesthesia and in the ICU ; spirometric and carbon dioxide monitoring in neonatal; pediatric and adult patients during general anesthesia and in the ICU and emergency department. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1286-2010 | 100-120V Capnogard 1265 - 100-120V Capnogard 1265; Refurbished; Respironics/Philips product.Intended to be used for monitoring end tidal CO2 and respiration rate in all critical monitoring environments in ventilatory support; patient support and anesthesia. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1276-2010 | TidalWave SP 710 - TidalWave SP 710; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1262-2010 | NICO2 - NICO2; Model 7600 Respironics/Philips Product.Intended use is cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1293-2010 | 100-200V NICO; 7300 Cardio Mgt System - 100-200V NICO; 7300 Cardio Mgt System; Refurbished; Respironics/Philips product.The intended use is cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation during general anesthesia and in the ICU ; spirometric and carbon dioxide monitoring in neonatal; pediatric and adult patients during general anesthesia and in the ICU and emergency department. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1285-2010 | NIC02 Model 7600 - NIC02 Model 7600 Refurbished; Respironics/Philips product.Intended use is cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation during general anesthesia and in the ICU ; spirometric and carbon dioxide monitoring in neonatal; pediatric and adult patients during general anesthesia and in the ICU and emergency department. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1305-2010 | 100-120V Oxypleth - 520A Pulse Oximeter - 100-120V Oxypleth - 520A Pulse Oximeter; Refurbished - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1264-2010 | C5 LOFLO Evaluation Kit - C5 LOFLO Evaluation Kit; Respironics/Philips product.Intended use of the LoFlo C5 CO2 sensor is to provide carbon dioxide monitoring. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1273-2010 | Tidal wave - Tidal wave; Model 610; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1261-2010 | Capnostat 5 Evaluation Kit - Evaluation Kit; Capnostat 5; Respironics/Philips Product.Intended use of the Capnostat 5 CO2 sensor is to provide carbon dioxide monitoring to a host monitoring system. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1265-2010 | LOFLO Engine Evaluation Kit - Evaluation Kit; LOFLO Engine; Respironics/Philips product.Intended use of the LoFlo C5 CO2 sensor is to provide carbon dioxide monitoring. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1270-2010 | Transcutaneous C02 & O2 Monitor - Transcutaneous C02 & O2 Monitor - 860; Respironics/Philips product.Used for trending O2 and CO2 tension at the skin surface for neonates and adults as an adjunct to arterial O2 and CO2 measurements. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1289-2010 | Tidal Wave 610 - Tidal Wave 610; Refurbished; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1300-2010 | 100-120V 515B with RS-232 Pulse Oximeter - 100-120V 515B with RS-232 Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1259-2010 | HeartStart MRx- AC power module kit - HeartStart MRx- AC power module kit.For use for the termination of ventricular tachycardia and ventricular fibrillation. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1263-2010 | Sharn TidalGuard SP 710V - Sharn TidalGuard SP 710V; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1290-2010 | 8100 Cosmo+ Resp Profile Monitor - Resp Profile Monitor; 8100 Cosmo+; Refurbished; Respironics/Philips product.Used to provide spirometric; carbon dioxide and pulse oximetry monitoring during general anesthesia and in the intensive care unit (ICU) and the emergency department (ED). | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1271-2010 | French 100-120V 860 TC C02 & 02 Monitor - French 100-120V 860 TC C02 & 02 Monitor; Respironics/Philips product.Used for trending O2 and CO2 tension at the skin surface for neonates and adults as an adjunct to arterial O2 and CO2 measurements. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1257-2010 | HeartStart MRx Defibrillator/Monitor - AC power cords manufactured by Electri-Cord with tall bridge design. Used across all Philips product lines; as well as Respironics/Philips products; Dixtel owned products and Dixtal (formerly Novametrix or Respironics/Novametrix products. Philips HeartStart MRx Defibrillator/Monitor model number M3535A; ID number 861288.For use for the termination of ventricular tachycardia and ventricular fibrillation. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1267-2010 | Spanish Capnogard 1265 - Spanish Capnogard 1265; Respironics/Philips product.Intended to be used for monitoring end tidal CO2 and respiration rate in all critical monitoring environments. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1299-2010 | French 100-120V 515B Pulse Oximeter - French 100-120V 515B Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1309-2010 | 100-120V 2001 Pulse Oximeter with Mars Nurse Call - 100-120V 2001 Pulse Oximeter with Mars Nurse Call; Refurbished - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1287-2010 | Cosmo Model 7100 SP02 & ETC02 Monitor - Cosmo Model 7100 SP02 & ETC02 Monitor; Refurbished; Respironics/Philips product.Intended to be used for monitoring end tidal CO2; respiration rate; oxygen saturation and pulse rate in critical monitoring environments including ventilatory support and anesthesia. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1279-2010 | Japanese Tidalwave SP 715 - Japanese TidalWave SP 715; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1304-2010 | Spanish 100-120V 2001 Pulse Oximeter WI - Spanish 100-120V 2001 Pulse Oximeter WI - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1280-2010 | 100-120V NICO - 100-120V NICO; 7300 Cardio Mgt system; Respironics/Philips product.Intended use is cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation during general anesthesia and in the ICU ; spirometric and carbon dioxide monitoring in neonatal; pediatric and adult patients during general anesthesia and in the ICU and emergency department. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1307-2010 | 515B Pulse Oximeter - 515B Pulse Oximeter; RS232; Refurbished - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1308-2010 | 100-120V 2001 Pulse Oximeter with Mars - 100-120V 2001 Pulse Oximeter with Mars; Refurbished - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1284-2010 | NIC02 Model 7600 - NIC02 Model 7600 Demo Unit; Respironics/Philips product.Intended use is cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation during general anesthesia and in the ICU ; spirometric and carbon dioxide monitoring in neonatal; pediatric and adult patients during general anesthesia and in the ICU and emergency department. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1288-2010 | Transcutaneous C02 & 02 Monitor - Transcutaneous C02 & 02 Monitor - 860; Refurbished; Respironics/Philips product.Used for trending O2 and CO2 tension at the skin surface for neonates and adults as an adjunct to arterial O2 and CO2 measurements. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1291-2010 | TIDALWAVE SP 710 - REFURBISHED TIDALWAVE SP 710; REFURBISHED; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1295-2010 | 100-120V Oxypleth - 520A Pulse Oximeter - 100-120V Oxypleth - 520A Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1297-2010 | Span 100-120V 520A - OxyPleth Pulse Oximeter - Span 100-120V 520A - OxyPleth Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1302-2010 | 100-120V 2001 Pulse Oximeter With Mars - 100-120V 2001 Pulse Oximeter With Mars - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1292-2010 | Tidal Wave SP 715 - Refurbished Tidal Wave SP 715; Refurbished; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1298-2010 | 515B Pulse Oximeter - 515B Pulse Oximeter - Dixtal owned product.Intended use is to provide continuous; non-invasive monitoring of functional saturation (SpO2) and pulse rate. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1268-2010 | Cosmo - Model 7100 SP02 & ETC02 Monitor - Cosmo - Model 7100 SP02 & ETC02 Monitor; Respironics./Philips product.Intended to be used for monitoring end tidal CO2; respiration rate; oxygen saturation and pulse rate in critical monitoring environments. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1275-2010 | Spanish 100-120V Cosmo Plus 8100 - Spanish 100-120V Cosmo Plus 8100; Respironics/Philips product.Used to provide spirometric; carbon dioxide and pulse oximetry monitoring. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1294-2010 | Tidal Wave - 615 - Tidal Wave - 615; Refurbished; Respironics/Philips product.Intended use of the monitors is to provide short term monitoring of carbon dioxide and oxygen saturation during anesthesia/recovery; in the intensive care unit (ICU); and in Emergency Medicine/transport of Respiratory Care. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1281-2010 | French 100-120V NICO 7300 - French 100-120V NICO 7300; Respironics/Philips product.Intended use is: cardiac output monitoring via the method of partial rebreathing in adult patients receiving mechanical ventilation during general anesthesia and in the ICU ; spirometric and carbon dioxide monitoring in neonatal; pediatric and adult patients during general anesthesia and in the ICU and emergency department. | 2 | Apr-08-2010 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock.
Z-1235-2010 | CADD-Legacy¿ 1 Model 6400 Ambulatory Infusion Pump - CADD-LEGACY¿ 1 Model 6400; Ambulatory Infusion Pump; REF 21-6400-51; Rx Only; Smiths Medical MD; Inc..; St. Paul; MN 55112. | 2 | Apr-05-2010 | Smiths Medical ASD; Inc. | The CADD-Legacy? 1 Model 6400 Ambulatory Infusion Pump (Serial Number 370969) was shipped without documentation that the required accuracy testing had been performed during the manufacturing process. If the subject infusion pump did not go through the accuracy testing; and if it does not meet accuracy specifications; there is a potential for over or under-delivery.
Z-1075-2010 | Medical Action Industries; dialysis kit - Medical Action Industries; Dialysis Kit; using BD Q-Syte Luer Access Split Septum (BD #385100); Part number: 57977. | 1 | Mar-30-2010 | Medical Action Industries Inc | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-1074-2010 | Medical Action Industries; Venous Access Tray - Medical Action Industries; Venous Access Tray; using BD Q-Syte Luer Access Split Septum (BD #385100)); Part number: 69498. | 1 | Mar-30-2010 | Medical Action Industries Inc | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-1073-2010 | Medical Action Industries; IV Start Bundle Kit - Medical Action Industries; IV Start Bundle Kit; using MPS Acacia with BD Q-Syte (BD #385151); Part number: 69946B. | 1 | Mar-30-2010 | Medical Action Industries Inc | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-1080-2010 | Acacia IV Extension Set with BD Q-Syte - Acacia Catalog # 385166 Dual "Y" IV Extension Set 50cm Macro Bore; Spin Nut w/BD Q-Syte | 1 | Mar-30-2010 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly; possibly resulting in an air embolism.
Z-1079-2010 | Acacia IV Extension Set with BD Q-Syte - Acacia Catalog # 385165 "Y" IV Extension Set 20cm Macro Bore; Spin Nut w/BD Q-Syte | 1 | Mar-30-2010 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly; possibly resulting in an air embolism.
Z-1076-2010 | Acacia IV Extension Set with BD Q-Syte - Acacia Catalog # 385150 IV Extension Set 15cm Small Bore; Spin Nut w/BD Q-Syte | 1 | Mar-30-2010 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly; possibly resulting in an air embolism.
Z-1077-2010 | Acacia IV Extension Set with BD Q-Syte - Acacia Catalog # 385151 IV Extension Set 15cm Small Bore; Spin Nut w/BD Q-Syte | 1 | Mar-30-2010 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly; possibly resulting in an air embolism.
Z-1078-2010 | Acacia IV Extension Set with BD Q-Syte - Acacia Catalog # 385164 Tri-Extension IV Set 15cm Small Bore; Spin Nut w/BD Q-Syte | 1 | Mar-30-2010 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly; possibly resulting in an air embolism.
Z-1186-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 11mm x 20cm x 135¿; Product Part Number: 1038-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1191-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 10mm x 20cm x 135¿; Product Part Number: 1042-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1184-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 11mm x 20cm x 130¿; Product Part Number: 1036-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1192-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 12mm x 20cm x 135¿; Product Part Number: 1043-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-118-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 13mm x 20cm x 130¿; Product Part Number: 1037-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1182-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 9mm x 20cm x 130¿; Product Part Number: 1034-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1190-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 12mm x 20cm x 130¿; Product Part Number: 1041-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1189-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 12mm x 17cm x 130¿; Product Part Number: 1041-170. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1181-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 9mm x 17cm x 130¿; Product Part Number: 1034-170. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1183-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 11mm x 17cm x 130¿; Product Part Number: 1036-170. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1187-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 10mm x 17cm x 130¿; Product Part Number: 1040-170 | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1188-2010 | AOS Trochanteric Nail - AOS Trochanteric Nail; 10mm x 20cm x 130¿; Product Part Number: 1040-200. | 2 | Mar-22-2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail; where one piece was found to be incorrectly bent. If this issue were to recur; surgery time could be extended; but only slightly as nail exchanges take place commonly in a surgery of this type.
Z-1153-2010 | XACT Carotid Stent System - Xact 10 mm x 20 mm x 136 cm (Part Number 82099-01)The Xact Carotid Stent System; used in conjunction with Emboshield Embolic Protection System is indicated for the improvement of the lumen diameter of carotid arteries in patients considered at high risk for adverse events from carotid endarterectomy who require percutaneous carotid angioplasty and stenting for occlusive artery disease. | 2 | Mar-16-2010 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the XACT Carotid Stent System because the product was distributed prior to approval of a PMA supplement for a manufacturing line move between sites. There are no product quality issues associated with this action.
Z-1152-2010 | XACT Carotid Stent System - Xact 9 mm x 20 mm x 136 cm (Part Number 82089-01)The Xact Carotid Stent System; used in conjunction with Emboshield Embolic Protection System is indicated for the improvement of the lumen diameter of carotid arteries in patients considered at high risk for adverse events from carotid endarterectomy who require percutaneous carotid angioplasty and stenting for occlusive artery disease. | 2 | Mar-16-2010 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the XACT Carotid Stent System because the product was distributed prior to approval of a PMA supplement for a manufacturing line move between sites. There are no product quality issues associated with this action.
Z-1041-2010 | Pointe Scientific Liquid Hitachi Calcium Reagent Set - Calcium Reagent Set. Pointe Scientific Liquid Hitachi Calcium Reagent Set; Pointe Scientific; Inc; Canton; MI; Catalog #'s HC902-576 and HC902-1000. Calcium Reagent Set for repackaging; Catalog # HC902-SAM-911. The product is used for the quantitative determination of calcium in serum on Hitachi analyzers. | 2 | Mar-10-2010 | Pointe Scientific; Inc. | Non-compliance with the Good Manufacturing Practice regulations; as the manufacturing process for this product had not been validated.
Z-1043-2010 | Pointe Scientific Liquid Hitachi Calcium Reagent Set - Calcium (CPC) R2 Reagent Set. Pointe Scientific Liquid Calcium (CPC) R2 Reagent Set; Pointe Scientific; Inc; Canton; MI; Catalog #'s HC402-R2. Calcium CPC Hitachi Reagent Set for repackaging; Catalog # 3-HC902-R2. The product is used for the quantitative determination of calcium in serum on Hitachi analyzers. | 2 | Mar-10-2010 | Pointe Scientific; Inc. | Non-compliance with the Good Manufacturing Practice regulations; as the manufacturing process for this product had not been validated.
Z-1042-2010 | Pointe Scientific Liquid Hitachi Calcium Reagent Set - Calcium (CPC) R1 Reagent Set. Pointe Scientific Liquid Calcium (CPC) R1 Reagent Set; Pointe Scientific; Inc; Canton; MI; Catalog #'s HC402-R1. Calcium CPC Hitachi Reagent Set for repackaging; Catalog # 3-HC902-R1. The product is used for the quantitative determination of calcium in serum on Hitachi analyzers. | 2 | Mar-10-2010 | Pointe Scientific; Inc. | Non-compliance with the Good Manufacturing Practice regulations; as the manufacturing process for this product had not been validated.
Z-0919-2010 | BCI 3180 Pulse Oximeter - BCI 3180 Pulse Oximeter uses power cord WW3005 (Electri-Cord Manufacturing Co.; 312 E. Main St.; Westfield; PA 16950). | 2 | Mar-02-2010 | Smiths Medical PM; Inc. | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords; manufactured by Electri-Cord Manufacturing Co.; for use with BCI 3180 Pulse Oximeter(Human); 9200 Advisor Vital Signs Monitor(Human or Veterinary); 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq
Z-0920-2010 | 8400 Capnocheck Capnometer - 8400 Capnocheck Capnometer uses power cord 8404 (Electri-Cord Manufacturing; 312 E. Main St; Westfield; PA 16950) for its battery charger 330. | 2 | Mar-02-2010 | Smiths Medical PM; Inc. | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords; manufactured by Electri-Cord Manufacturing Co.; for use with BCI 3180 Pulse Oximeter(Human); 9200 Advisor Vital Signs Monitor(Human or Veterinary); 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq
Z-0918-2010 | 9200 Advisor Vital Signs Monitor - 9200 Advisor Vital Signs Monitor uses power cord WW3005 (Electri-Cord Manufacturing Co.; 312 Main St.; Westfield; PA 16950). | 2 | Mar-02-2010 | Smiths Medical PM; Inc. | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords; manufactured by Electri-Cord Manufacturing Co.; for use with BCI 3180 Pulse Oximeter(Human); 9200 Advisor Vital Signs Monitor(Human or Veterinary); 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq
Z-0914-2010 | Access Peristaltic Pump - Access Peristaltic Pump component of: UniCel DxI 800 Access Immunoassay Systems; UniCel DxI 600 Access Immunoassay Systems; UniCel DxC 880i SYNCHRON Access Clinical Systems.Intended to be used with the Access Immunoassay System to remove waste from the system. | 2 | Mar-02-2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of a new premature failure of the upper aspirate peristaltic pump in the systems listed above. Premature failure of the pump may negatively affect precision. The failure is limited to those UniCel DXI 800; DXI 600; and DxC 880i systems which have a pump identified with one of the new pump manufacturing numbers (Mfg. Nos.) listed ab
Z-0751-2010 | BD Q-Syte Extension Set; Luer access extension - BD Q-Syte Extension Set; 15cm Std Bore; P.V. 1.14 ml Luer-Lok; REF 385101; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0753-2010 | BD Q-Syte Vial Access Adapter; Luer access adapter - BD Q-Syte Vial Access Adapter; 0.16 ml; REF 385108; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0830-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383531; 24GA 0.75IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0750-2010 | BD Q-Syte; Luer Access connector - BD Q-Syte; Luer Access Split-Septum; 0.10 ml; REF 385100; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0834-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383536; 20GA 1.00IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0832-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383533; 20GA 1.00IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0836-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383538; 20GA 1.75IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0755-2010 | MPS Acacia Extension Set; Luer access extension - MPS Acacia Extension Set; 15cm Small Bore; Spin Nut w/BD Q-Syte; REF 385151 | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0752-2010 | BD Q-Syte Extension Set; Luer access extension - BD Q-Syte Extension Set; 15 cm; P.V. 0.21 ml; REF 385102; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0831-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383532; 22GA 1.00IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0838-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383540; 18GA 1.75IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0756-2010 | MPS Acacia Tri-Extension Set; Luer access extension - MPS Acacia Tri-Extension Set; 15cm Small Bore; Spin Nut w/BD Q-Syte; REF 385164 | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0839-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383647; 20GA 1.25IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0840-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383649; 18GA 1.25IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0837-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383539; 18GA 1.25IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0829-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383530; 24GA 0.56IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0833-2010 | Closed IV Catheter system - BD Nexiva Closed IV Catheter System; REF 383534; 20GA 1.25IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0835-2010 | Closed IV Catheter system - BD Nexiva HF Closed IV Catheter System; REF 383537; 20GA 1.25IN; Becton Dickinson Infusion Therapy Systems Inc. | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0754-2010 | MPS Acacia Extension Set; Luer access extension - MPS Acacia Extension Set; 15cm RightBore-18; Spin Nut w/BBD Q-Syte; REF 385150; | 1 | Feb-02-2010 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter; possibly resulting in an air embolism.
Z-0757-2010 | TRIATHLON TS KNEE SYSTEM - Triathlon Total Stabilizer Femoral Component; Size # 2for Use with TS Components; Sterile; Made in the USAStryker Orthopaedics; Mahwah; NJ 07430.The Triathlon Total Knee System components are intended for use in primary and revision total knee arthroplasty to alleviate pain and restore function. | 2 | Feb-01-2010 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing; possibly preventing the assembly of a stem extension.
Z-0760-2010 | TRIATHLON TS KNEE SYSTEM - Triathlon Total Stabilizer Femoral Component; For Use with TS Components; Sterile; Made in the USAStryker Orthopaedics; Mahwah; NJ 07430.The Triathlon Total Knee System components are intended for use in primary and revision total knee arthroplasty to alleviate pain and restore function. | 2 | Feb-01-2010 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing; possibly preventing the assembly of a stem extension.
Z-0758-2010 | TRIATHLON TS KNEE SYSTEM - Triathlon Total Stabilizer Femoral Component; For Use with TS Components; Sterile; Made in the USAStryker Orthopaedics; Mahwah; NJ 07430.The Triathlon Total Knee System components are intended for use in primary and revision total knee arthroplasty to alleviate pain and restore function. | 2 | Feb-01-2010 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing; possibly preventing the assembly of a stem extension.
Z-0761-2010 | TRIATHLON TS KNEE SYSTEM - Triathlon Total Stabilizer Femoral Component; For Use with TS Components; Sterile; Made in the USAStryker Orthopaedics; Mahwah; NJ 07430.The Triathlon Total Knee System components are intended for use in primary and revision total knee arthroplasty to alleviate pain and restore function. | 2 | Feb-01-2010 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing; possibly preventing the assembly of a stem extension.
Z-0759-2010 | TRIATHLON TS KNEE SYSTEM - Triathlon Total Stabilizer Femoral Component; For Use with TS Components; Sterile; Made in the USAStryker Orthopaedics; Mahwah; NJ 07430.The Triathlon Total Knee System components are intended for use in primary and revision total knee arthroplasty to alleviate pain and restore function. | 2 | Feb-01-2010 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing; possibly preventing the assembly of a stem extension.
Z-0745-2010 | Cardiac Cath Pack Kit - Custom Cardiac Cath Pack Kit; K09-08945; Sterile EO; Merit Medical Systems; Inc.; South Jordan; Utah 84095.Intended for fluid management convenience when catheterization is required. | 2 | Jan-29-2010 | Merit Medical Systems; Inc. | A manufacturing defect may prevent adequate sterilization of the entire device.
Z-0746-2010 | Custom Angiographic Kit - Custom Angiographic Kit; K09-MS3084; Sterile EO; Merit Medical Systems; Inc.; South Jordan; Utah 84095.Intended for fluid management convenience. | 2 | Jan-29-2010 | Merit Medical Systems; Inc. | A manufacturing defect may prevent adequate sterilization of the entire device.
Z-0748-2010 | Cath Procedure Tray - Jeanes Hospital Custom Cath Procedure Tray; K09T-08945A; Sterile EO; Merit Medical Systems; Inc.; 1600 West Merit Parkway; South Jordan; Utah 84095.Intended for fluid management convenience when catheterization is required. | 2 | Jan-29-2010 | Merit Medical Systems; Inc. | A manufacturing defect may prevent adequate sterilization of the entire device.
Z-0747-2010 | Cath Procedure Tray - Jeanes Hospital Custom Cath Procedure Tray; K09T-08945; Sterile EO; Merit Medical Systems; Inc.; 1600 West Merit Parkway; South Jordan; Utah 84095.Intended for fluid management convenience when catheterization is required. | 2 | Jan-29-2010 | Merit Medical Systems; Inc. | A manufacturing defect may prevent adequate sterilization of the entire device.
Z-0744-2010 | Cardiac Cath Pack - University Health System Custom Cardiac Cath Pack; K09T-02178; Srterile EO; Merit Medical Systems; Inc.; 1600 West Merit Parkway; South Jordan; Utah 84095.Intended for fluid management convenience when catheterization is required. | 2 | Jan-29-2010 | Merit Medical Systems; Inc. | A manufacturing defect may prevent adequate sterilization of the entire device.
Z-0488-2010 | Elements Gutta Percha and Elements RealSeal Cartridges - Elements Gutta Percha and Elements RealSeal Cartridges; Part Number: 972-1002; 972-1003; 972-2001; 972-2002 | 2 | Dec-09-2009 | Sybronendo | The recall was initiated because during the manufacturing process of the affected units; the molding process resulted in lock nuts that did not have the level of acceptability required; and therefore; may not withstand the pressures exhibited by the Elements Obturation Unit with which the cartridges are used. More Specifically; there exists a possibility that if the lock nut were to fail; the cart
Z-0198-2010 | VerSys Hip System Femoral Head - Zimmer VerSys Hip System Femoral Head; 12/14 taper; 28 mm dia.; neck length +0 mm; sterile; Zimmer; Warsaw; IN; REF 8018-28-02.Intended for mating femoral stems equipped with tapered necks. | 2 | Nov-13-2009 | Zimmer Inc. | Manufacturing material may be present on the surface of the implant.
Z-0086-2010 | Pivot BiPolar Head - Ortho Development Corp.; Pivot BiPolar Head; 28mm; Hip Prosthesis; Femoral Head. | 2 | Oct-28-2009 | Ortho Development Corporation | A Pivot BiPolar Head; the femoral head in a hip prosthesis; might not move freely after assembly due a stacking up of manufacturing tolerances.
Z-1814-2009 | Engenex - ConvaTec; Engenex Tunnel Dressing; for use with Engenex Negative Pressure Wound Therapy; (1 x 15 cm)Rx only; Latex Free; SterileHealth Industry Bar Code: H984BWS4464509A.Packaged in boxes of 15 units per box. | 3 | Sep-29-2009 | ConvaTec | Labeling: the actual manufacturing date (2/2009) and actual expiration date (2/2011) are switched.
Z-2227-2009 | 3i Custom - 3i Custom # 5753; OSCM2913; 2.9MMX 3.4MM X 13MM; OSSEOTITE CUSTOM; MICROMINI IMPLANT; LOT # 470258; Sterile. Implant Innovations; Inc. 4555 Riverside Drive; Palm Beach Gardens; FL 33410.Custom temporary cover screws are for use in maintaining the soft tissue during the healing process of a dental restoration. | 2 | Sep-22-2009 | Biomet 3i; LLC | Healing Abutment was not made within required manufacturing specification. The product did not have the relief behind the thread; which does not allow proper seating.
Z-2049-2009 | Penumbra Neuron Select Catheter - Penumbra Neuron 5F Select Catheter; Neuro Deliver/Intracranial Access System Catheter; manufactured by Penumbra; Inc.; Alameda; CA.Intended for the introduction of interventional devices into the peripheral; coronary and neuro vasculature. | 1 | Sep-08-2009 | Penumbra Inc. | Leakage-- Due to a manufacturing error; some products may have a fold in the innermost layer; causing disruption in the layering process and may result in blow holds or exposed wire braids; which may result in further damage.
Z-2046-2009 | Terumo Weighted Flexible Sucker - Terumo Sarns cardiovascular disposable products; Weighted Flexible Sucker; sterile; Terumo Cardiovascular Systems Corp.; Ann Arbor; MI.; Catalog No. 804113. | 2 | Aug-04-2009 | Terumo Cardiovascular Systems Corporation | A loose strand of stainless steel left on the device during manufacturing may result in an unretrieved metal fragment remaining in the patient.
Z-1598-2009 | Fusion Omni-Tome - Fusion Omni-Tome; Rx Only; Wilson Cook Medical; GI Endoscopy; 4900 Bethania Station Road; Winston-Salem; NC 27105. | 3 | Jul-21-2009 | Cook Endoscopy | The wire control port of these devices may be missing due to omission of a manufacturing activity.
Z-1009-2009 | Roche COBAS c 111 Analyzer - Roche COBAS c 111 Analyzer with ISE module option; GMMI No. 04777433001; Roche Diagnostics; Indianapolis; IN.Is an in-vitro diagnostic analyzer capable of performing clinical chemistry; specific protein electrolyte tests. Analytes are measured photometrically or turbidimetrically. The analyzer also has an optional ISE module for measuring sodium; potassium and chloride. | 2 | Mar-24-2009 | Roche Diagnostics Corp. | Due to a manufacturing error at the supplier; not all cobas c 111 pinch valve caps are properly glued to the correct position on the pinch valve shaft. Without the glue; the cap can become loose or misaligned. The effect of these faulty valves is that they may not open or close properly which can disturb the fluid transportation within the ISE module resulting in problems with ISE calibration; con
Z-1033-2009 | InteliFUSE; Drill Bit; 2.00 mm; J-Latch - InteliFUSE; Drill Bit; 2.00 mm; J-Latch; QTY: 1; REF: DB.200.JL; non-sterile; 1515 Poydras st.; Ste 1490; New Orleans; LA 70112; MDCI Arundel House; 1 Liverpool Gardens; Worthing; West Sussex BN11 1SL; United KingdomIntended Use: Creating of holes for implant placement. | 2 | Mar-05-2009 | Intelifuse Inc | Manufacturing specification for drill bit was incorrect which; when used in conjunction with the drill guide; could cause the production of metal shavings.
Z-1032-2009 | InteliFUSE; Drill Bit; 1.50 mm; J-Latch - InteliFUSE; Drill Bit; 1.50 mm; J-Latch; QTY: 1; REF: DB.150.JL; non-sterile; 1515 Poydras st.; Ste 1490; New Orleans; LA 70112; MDCI Arundel House; 1 Liverpool Gardens; Worthing; West Sussex BN11 1SL; United KingdomIntended Use: Creating of holes for implant placement. | 2 | Mar-05-2009 | Intelifuse Inc | Manufacturing specification for drill bit was incorrect which; when used in conjunction with the drill guide; could cause the production of metal shavings.
Z-0992-2009 | Bone Wax - Bone Wax. Ethicon; Inc.Bone wax is available sterile in individual foil envelopes; each containing 2.5 grams; and then packaged in an individually sealed over wrap packet.The device is indicated for use to control bleeding from the bone surface during surgical procedures. | 2 | Mar-04-2009 | Ethicon; Inc. | A hospital reported to their Ethicon sales representative that they had an "unmarked" box of bone was on their shipping dock that they could not receive due to lack of appropriate documentation. Upon investigation; this box was determined to be missing from a shipment between two steps of the manufacturing process. Further checks determined that two other boxes were missing.
Z-0508-2009 | Novel VBR Spinal System - Novel PEEK VBR XS 16-22 MM; S-5; 26MM; Part Numbers : 64753-126. Indicated for use in the thoracolumbar spine (T1 to L5) for partial or total replacement of a collapsed; damaged or unstable vertebral body due to tumor or trauma (i.e.; fracture). Intended for use with supplemental spinal fixation systems. | 2 | Dec-31-2008 | Alphatec Spine; Inc. | Firm confirmed that two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number & lot number. They were incorrectly laser marked with the wrong part numbers and sizes. Two lots were mixed from the following lot numbers: Part Number 64733-120 Lot Number 611884; and Part Number 64753-126 and Lot Number 611881.
Z-0507-2009 | Novel VBR Spinal System - Novel PEEK VBR XS 16-22 MM; S-5; 20MM;Part Number: 64733-120. Indicated for use in the thoracolumbar spine (T1 to L5) for partial or total replacement of a collapsed; damaged or unstable vertebral body due to tumor or trauma (i.e.; fracture). Intended for use with supplemental spinal fixation systems. | 2 | Dec-31-2008 | Alphatec Spine; Inc. | Firm confirmed that two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number & lot number. They were incorrectly laser marked with the wrong part numbers and sizes. Two lots were mixed from the following lot numbers: Part Number 64733-120 Lot Number 611884; and Part Number 64753-126 and Lot Number 611881.
Z-0153-2009 | Allez Spine Laguna Pedicle Screw System - Allez Spine Laguna Pedicle Screw System; Catalog Number: IT-FN1003Del Mar Pedicle Screw System; IT-ST1001Vertebrae Rotation Tool;IT-H W1001Hex Wrench; IT-CB1001Coronal Bender-L; IT-CB1002 Coronal Bender-R; IT-PH1001Power Rod Holder; IT-BP1001 Thoracic Ball Handle Probe; Curved; IT-BP1002 Thoracic Ball Handle Probe; Straight; DM-MR0530Monoaxial Screw Non-Winged; Size 5 x 30; DM-MR0535 Monoaxial Screw Non-Winged; Size 5 x 35; DM-MR0540 Monoaxial Screw Non-Winged; Size 5 x 40; DM-MR0545 Monoaxial Screw Non-Winged; Size 5 x 45; DM-MW0530 Monoaxial Screw Winged; Size 5 x 30; DM-MW0535 Monoaxial Screw Winged; SizeS x 35; DM-MW0540 Monoaxial Screw Winged; Size S x 40; DM-MW0545Monoaxial Screw Winged; Size 5 x 45; DM-MR0630 Monoaxial Screw Non-Winged; Size 6 x 30; DM-MR0635 Monoaxial Screw Non-Winged; Size 6 x 35;DM-MR0640 Monoaxial Screw Non-Winged; Size 6 x 40; DM-MR0645 Monoaxial Screw Non-Winged; Size 6 x 45; DM-MR0650 Monoaxial Screw Non-Winged; Size 6 x 50; DM-MW0630 Monoaxial Screw Winged; Size 6 x 30;DM-MW0635 Monoaxial Screw Winged; Size 6 x 35; DM-MW0640Monoaxial Screw Winged; Size 6 x 40; DM-MW0645 Monoaxial Screw Winged; Size 6 x 45; DM-MR0735 Monoaxial Screw Non-Winged; Size 7 x 35DM-MR0740 Monoaxial Screw Non-Winged; Size 7 x 40; DM-MR0745 Monoaxial Screw Non-Winged; Size 7 x 45; RD-ST5541 Straight Rod; 4mm HexIntended to help provide correction; immobilization and stabilization of spinal segments as an adjunct to fusion of the thoracic; lumbar and/or sacral space. | 2 | Oct-28-2008 | Allez Spine; LLC | This action is being taken voluntarily as a result of a Warning Letter issued to AlIez Spine LLC by the Food & Drug Administration. The Company decided it could better address the FDA's concerns regarding Current Good Manufacturing Practice requirements of the Quality System without any product in the market so it could focus all of its efforts on responding to the FDA.
Z-0155-2009 | Allez Spine Cross Connectors - Allez Spine Cross Connectors; Catalog Number: LG-CC1001Cross Link; Small 30-37mm; LG-CC1002 Cross Link; Medium 37-50mm; LG-CC1003Cross Link; Lange 50-80mm; lT-T01003 Cross Link Torque Hex Driver; 2.5mm | 2 | Oct-28-2008 | Allez Spine; LLC | This action is being taken voluntarily as a result of a Warning Letter issued to AlIez Spine LLC by the Food & Drug Administration. The Company decided it could better address the FDA's concerns regarding Current Good Manufacturing Practice requirements of the Quality System without any product in the market so it could focus all of its efforts on responding to the FDA.
Z-0154-2009 | Allez Spine Del Mar Pedicle Screw System - Allez Spine Del Mar Pedicle Screw System; Catalog Number: IT-FN1001Laguna Pedicle Screw System Tray #1; IT-FN 1002 Laguna Pedicle Screw System Tray #2; lT-CT1001Counter Torque (w/Large Rubber Handle); IT-HP1001 Head Positioner (w/Rubber Handle); IT-RP1001Rod Pusher;lT-RK1001 Rocker (Goal Post Style Tip); lT-BA1002 Vampire Awl; IT-RG1001Rod Gripper; Ratcheting (Long Nose); IT-TP1001Tap; 4mm (Thin Style; Gold Tip); IT-TP1002 Tap; 5mm (Thin Style; Gold Tip) IT-TP1003 Tap; 6mm (Thin Style; Gold Tip); IT-TP1004 Tap; 7mm (Thin Style; Gold Tip); IT-SD 1002 Multiaxial Screwdriver (No Windows); IT-SD1005 Screwdriver Shaft 75 (w/BaIl Bearing); IT-CT1002 Counter Torque (Short No Windows); IT-SD1006 Multiaxial Screwdriver (w/Windows);IT-TP1005 Tap; 4mm (Silver Tn-Fluted);IT-TP1006 Tap; 5mm (Silver Tn-Fluted);IT-TP1007 Tap; 6mm (Silver Tn-Fluted);IT-TP1008 Tap; 7mm (Silver Tn-Fluted); lT-HP1002 Head Positioner (All Metal Handle);IT-RK1002 Rocker (Horse Shoe Style Tip);IT-SD1007 Screwdriver Shaft 8.5 (w/Ball Bearing);IT-CT1003 Counter Torque (Short w/Windows);IT-TP1014 Tap; 4mm (Silver Circumferential);IT-TP1015 Tap; 5mm (Silver Circumferential);IT-TP1016 Tap; 6mm (Silver Circumferential);IT-TP1017 Tap; 7mm (Silver Circumferential);IT-SD1007 Screwdriver Shaft 8.5 (No Ball Bearing);IT-XR1001 Extension Remover;IT-S01006 Screwdriver_Shaft 10 (No Ball Bearing);IT-FB1001 French Rod Bender;IT-SP1001 Ball Tip Probe;IT-TR1001 Trial Rod;IT-RG 1002 Rod Gripper; Ratcheting (Short Nose);IT-BP1003 Lumbar Probe; Straight (Long);IT-BP1004 Lumbar Probe; Curved (Long);IT-SP1002 Sounder (Ball Tip) Probe; Heavy Duty;RD-ST5540 Rod 5;5 x 400mm; Straight;IT-BA1001 Newport Awl;IT-PU 1001 Persuader; IT-RR1001 Retaining Ring;lT-NS1001 In-Situ Left;IT-N51002 In-Situ Right;IT-BP100S Lumbar Probe; Straight (Short);IT-BP1006 Lumbar Probe; Curved (Short);IT-CP1001 Compressor;IT-DT1001 Distractor;IT-BP1001 Thoracic Ball Handle Probe; Curved;IT-BP1002 Thoracic Ball Handle Probe; Straight;IT-RH1002 Ratcheting T-Handle;IT-NC1001 Locking Nut Caddy;IT-SG1001 Screw Gauge Block;IT-RH1002 Ratcheting Handle; Straight;IT-T01002 Preset Torque Wrench (With Connector);LG-LN1002 Locking Nut; Single Piece;LG-PR0530C Polyaxial Screw; Non-Winged; 5 X 30;LG-PR0535C Polyaxial Screw; Non-Winged; 5 X 35;LG-PR0540C Polyaxial Screw; Non-Winged; 5 X 40;LG-PR0545C Polyaxial Screw; Non-Winged; 5 X 45;LG-PR0550C Polyaxial Screw; Non-Winged; 5 X 50;LG-PR0630C Polyaxial Screw; Non-Winged; 6 X 30;LG-PR0635C Polyaxial Screw; Non-Winged; 6 X 35;LG-PR0640C Polyaxial Screw; Non-Winged; 6 X 40;LG-PR0645C Polyaxial Screw; Non-Winged; 6 X 45;LG-PR0650C Polyaxial Screw; Non-Winged; 6 X 50;LG-PR0730C Polyaxial Screw; Non-Winged; 7 X 30;LG-PR0735C Polyaxial Screw; Non-Winged; 7 X 35;LG-PR0740C Polyaxial Screw; Non-Winged; 7 X 40;LG-PR074SC Polyaxial Screw; Non-Winged; 7 X 45;LG-PR0750C Polyaxial Screw; Non-Winged; 7 X 50;LG-PR0755C Polyaxial Screw; Non-Winged; 7 X 55;LG-PR083SC Polyaxial Screw; Non-Winged; 8 X 35;LG-PR0840C Polyaxial Screw; Non-Winged; 8 X 40;LG-PR0845C Polyaxial Screw; Non-Winged; 8 X 45;LG-PR0850C Polyaxial Screw; Non-Winged; 8 X 50;LG-PR0855C Polyaxial Screw; Non-Winged; 8)( 55;LG-PW0535C Polyaxial Screw; Winged; S X 35;LG-PW0540C Polyaxial Screw; Winged; 5 X 40;LG-PW0545C Polyaxial Screw; Winged; 5 X 45;LG-PW0635C Polyaxial Screw; Winged; 6 X 35;LG-PW0640C Polyaxial Screw; Winged; 6 X 40;LG-PW0645C Polyaxial Screw; Winged; 6 X 45;LG-PW0735C Polyaxial Screw; Winged; 7 X 35;LG-PW0740C Polyaxial Screw; Winged; 7 X 40;LG-PW0745C Polyaxial Screw; Winged; 7 X 45;RD-PB5504 ROD 5.5 X 40; PRE-BENT;RD-PB5505 ROD 5.5 X 50; PRE-BENT;RD-PB5506 ROD 5.5 X 60; PRE-BENT;RD-PB5507 ROD 5.5 X 70; PRE-BENT;RD-PB5508 ROD 5.5 X 80; PRE-BENT;RD-PB5509 ROD 5.5 X 90; PRE-BENT;RD-PB5510 ROD 5.5 X 100; PRE-BENT;RD-PB5511 ROD 5.5 X 110; PRE-BENT;RD-PB5 | 2 | Oct-28-2008 | Allez Spine; LLC | This action is being taken voluntarily as a result of a Warning Letter issued to AlIez Spine LLC by the Food & Drug Administration. The Company decided it could better address the FDA's concerns regarding Current Good Manufacturing Practice requirements of the Quality System without any product in the market so it could focus all of its efforts on responding to the FDA.
Z-2233-2008 | Hamilton Soft Grip Pipettes; 1 ml adjustable volume - Hamilton Soft Grip Pipettes; 1 ml adjustable volume; model number 55019-40Product is indicated for manual pipetting. | 2 | Sep-17-2008 | Hamilton Co | Manufacturing issue with spring may restrict product to drawing less than 1/2 of its stated measured capacity
Z-2232-2008 | Hamilton Soft Grip Pipettes; 1 ml fixed volume - Hamilton Soft Grip Pipettes; 1 ml fixed volume; model number 55019-19Product is indicated for manual pipetting. | 2 | Sep-17-2008 | Hamilton Co | Manufacturing issue with spring may restrict product to drawing less than 1/2 of its stated measured capacity
Z-1183-2008 | Access Power Supply Assembly Sled - Access Power Supply Assembly SledPart Number: A48196/6845C; Component of: Access and Access 2 Immunoassay Systems; SYNCHRON LXi 725 System; and UniCel DxC 600i System; for in vitro diagnostic use. | 2 | Aug-20-2008 | Beckman Coulter Inc | Electrical grounding failure: During manufacturing; internal testing found one instrument wlth an intermittent failure ofthe electrical grounding. The source of the problem was traced to inadequate soldering.
Z-1191-2008 | Integra NeuroSensor - INTEGRA NeuroSensor Probe; REF NS-P; component of Cerebral Blood Flow and Intracranial Pressure Monitoring Kit; Integra NeuroSciences | 2 | Aug-05-2008 | Integra Neuro Sciences | Incorrect pressure reading: Due to a manufacturing error; the ICP reading could be inaccurate to the extent that they exceed label claims. Specifically; the Insertion Error label claim of +/- 3mmHg could be exceeded by up to -3mmHg; that could result in a maximum Insertion Error of +/- 6mmHg.
Z-1556-2008 | CORDIS S.M.A.R.T.¿ Control" Nitinol Stent System - CORDIS S.M.A.R.T.¿ Control" Nitinol Stent SystemModel Number(s): C10040SL; Manufactured in Mexico.Indicated for improving luminal diameter in patients with symptomatic atherosclerotic disease. | 2 | Jul-31-2008 | Cordis Corporation | Data transcription error. The 10 affected lots of Catalog Number C10040SL were manufactured with a prior version of the guidewire lumen material; which is not validated with the current manufacturing process. This condition may result in a separation of the device. If a separation was to occur; and the separated fragment of lumen material was not retrieved on the guidewire; blockage of the vessel
Z-1173-2008 | Trident - Trident PSL Acetabular Shells; Trident PSL HA Solid Back;40 mm to 73 mm; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1168-2008 | Trident - Trident Hemispherical Solid BK Acetabular Shells;42mm to 74mm; Hydroxylapatite Coated; Multiholed; Use with Size A insert; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1169-2008 | Trident - Trident Hemispherical Acetabular Shells HA Solid BK42mm- 74mm;Arc Deposited;Hydroxylapatite Coated; Multiholed; Use with Size A insert; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1174-2008 | Trident - Trident PSL HA Cluster:40 mm to 72 mm; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1171-2008 | Trident - Trident Hemispherical HA Cluster; 42 mm - 74mmArc Deposited;Hydroxylapatite Coated; Multiholed; Use with Size A insert; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1170-2008 | Trident - Trident Hemispherical Cluster;42mm - 74 mmArc Deposited;Hydroxylapatite Coated; Multiholed; Use with Size A insert; hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1172-2008 | Trident - Trident Hemispherical Multi Arc Deposited;Hydroxylapatite Coated; Multiholed; Use with Size A insert;hip prosthesis component; Stryker Orthopaedics; Howmedica Osteonics Corp; Mahwah; NJ 07430 | 2 | Jun-12-2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals.
Z-1132-2008 | Keeler All Pupil II Indirect Ophthalmoscope - Keeler All Pupil II Indirect Ophthalmoscope; Keeler Instruments; Inc.; Broomall; PA 19008 | 2 | May-22-2008 | Keeler Instruments Inc | Good Manufacturing Practices (GMP) deficiencies may compromise the safety and effectiveness of the device. Firm is on import alert 89-04 and released product.
Z-1311-2008 | Terumo "Custom" Cardiovascular Procedure Kit - Cardiovascular Procedure Kit; P/N 64789-01; Wisconsin Cardiovascular Services Pump and Table Pack; Terumo Cardiovascular; Ashland; MA. (the Pump & table pack contains a weighted flexible sucker; Catalog No. 804113) | 2 | Mar-28-2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient.
Z-1310-2008 | Terumo "Custom" Cardiovascular Procedure Kit - Cardiovascular Procedure Kit; P/N: 70932-01; HCSG-St. Mary's Hospital; Richmond; VA ; X-coated Pump and Table Pack; Terumo Cardiovascular; Ashland; MA (the X-coated pump and table pack contains a weighted flexible sucker; Catalog No.: 804113) | 2 | Mar-28-2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient.
Z-1313-2008 | Terumo "Custom" Cardiovascular Procedure Kit - Cardiovascular Procedure Kit; P/N 70579-04; St. Louis University Hospital Pump & Table Pack; Terumo Cardiovascular; Ashland; MA. (the Pump & Table pack contains a weighted flexible sucker; Catalog No. 804113) | 2 | Mar-28-2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient.
Z-1312-2008 | Terumo "Custom" Cardiovascular Procedure Kit - Cardiovascular Procedure Kit; P/N 70295-02; St. Louis University Hospital Pump & Table Pack; Terumo Cardiovascular; Ashland; MA. ( Pump & table pack contains a weighted flexible sucker; Catalog No. 804113) | 2 | Mar-28-2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient.
Z-0460-2008 | Alaris Pump Module (formerly Medley Pump Module) - Cardinal Health Alaris Pump Module (formerly Medley Pump Module); Model 8100; Cardinal Health; San Diego; CA. (All Alaris Pump modules shipped prior to December 23; 2007 are subject to this recall) | 1 | Jan-04-2008 | Cardinal Health 303 Inc DBA Alaris Products | Inaccurate flow rate; related to misassembled (missing; bent or broken) springs during the manufacturing or servicing of the mechanism assembly
Z-0363-2008 | Ascent POCTS (Posterior Occipital Cervical Thoracic System) - Ascent POCTS (Posterior Occipital Cervical Thoracic System) Part Numbers: 65-33 XX Series (65-3310;-3312; -3314; -3318; -3320-; -3322; -3324; -3326; -3328; and -3330); Multi-Axial Screw; 4.0mm x 26mm;Implant grade Titanium Alloy; Blacksone Medical Inc.; 90 Brookdale Drive; Springfield; MA 01104. | 2 | Dec-11-2007 | Blackstone Medical; Inc. | Manufacturing Defect: The saddle of the screw's body relative to the keyway exceeds the tolerance outlined in the device specifications. The extent to which the tolerance is exceeded varies from part to part; in the extreme the tolerance is exceeded by (+) 50%. As a consequence; there is a potential for set-screw cross-threading during implantation. Cross-threading could in turn lead to a comprom
Z-0116-2008 | Boston Scientific - Boston Scientific Radial Jaw Large Capacity Biopsy Forceps; Sterile; 240 mm. Box of 20U PN Number: M00513321REF: 1332-20 | 2 | Oct-31-2007 | Boston Scientific Corporation | Device lacked manufacturing finishing process resulting in endoscope damage.
Z-0117-2008 | Boston Scientific - Boston Scientific Radial Jaw Large Capacity Biopsy Forceps with Needle; Sterile; 240 mm. Box of 20UPN Number: M00513331REF: 1333-20 | 2 | Oct-31-2007 | Boston Scientific Corporation | Device lacked manufacturing finishing process resulting in endoscope damage.
Z-0118-2008 | Boston Scientific - Boston Scientific Radial Jaw Large Capacity Biopsy Forceps with Needle; Sterile; 240 mm. Box of 40UPN Number: M00513332REF: 1333-40 | 2 | Oct-31-2007 | Boston Scientific Corporation | Device lacked manufacturing finishing process resulting in endoscope damage.
Z-1207-2007 | TENS Device - Avazzia Tennant Biomodulator¿; Model: BEST AV-1; P/N #TB04303-103; manufactured by Avazzia; Inc.; Dallas; TX 75240. | 3 | Sep-01-2007 | Avazzia; Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications
Z-1209-2007 | TENS Device - Avazzia Med-Best Therapeutic Massager; Model: BEST AV-1; P/N #MB04182-103; manufactured by Avazzia; Inc.; Dallas; TX 75240. | 3 | Sep-01-2007 | Avazzia; Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications
Z-1210-2007 | TENS Device - Avazzia Med-Sport Massager; Model: BEST AV-1; P/N #MS04182-103; manufactured by Avazzia; Inc.; Dallas; TX 75240. | 3 | Sep-01-2007 | Avazzia; Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications
Z-1208-2007 | TENS Device - Avazzia BEST-RSI; Model: BEST AV-1; P/N #RS04303-103; manufactured by Avazzia; Inc.; Dallas; TX 75240. | 3 | Sep-01-2007 | Avazzia; Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications
Z-1042-2007 | OneTouch Ultra Test strip - OneTouch Ultra Test strip; a component of the following OneTouch Ultra Glucose monitoring systems: a)OneTouch Ultra; b) OneTouch InDuo; c) OneTouch Ultra2; d) OneTouch UltraMini; e) OneTouch UltraSmart. 25 strips per vial. Product comes in packs of 100 and 50. (Part numbers: 020-245-07; 020-244-07; 020-963-01) | 2 | Jul-13-2007 | Lifescan Inc | Damaged Vials/Inaccurate test results: a manufacturing equipment problem on one of the packaging lines had the potential to puncture vials; causing a single hole in the side of the vial below the rim. This result of this defect may produce inaccurately high or low blood glucose test results or error message readings.
Z-0829-2007 | Endura No-React - EnDura No-React Dural Substitute 5 cm x 6 cm | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0831-2007 | Endura No-React - EnDura No React Dural Substitute 10 cm x 12 cm | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0830-2007 | Endura No-React - Endura No-React Dural Substitute 6 cm x 10 cm | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0827-2007 | Endura No-React - Endura No-React Dural Substitute 2 cm x 10 cm | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0828-2007 | Endura No-React - EnDura No-React Dural Substitute 4 cm X 4 cm. | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0832-2007 | Endura No-React - EnDura No-React Dural Substitute 12 cm Diameter | 1 | May-24-2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences; the distributor of EnDura No-React Dural Substitute; manufactured by Shelhigh; Inc.; Union; NJ; is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t
Z-0692-2007 | BioHorizons 0.035 Hex Driver - BioHorizons 0.035 Hex Driver; REF 300-377; BioHorizons Implant Systems; Inc.; Birmingham; AL 35243; non-sterile; Rx only; REF 300-377. | 3 | Apr-13-2007 | BioHorizons Implant Systems Inc | Premature wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear.
Z-1164-2007 | BioHorizons 0.050 Handpiece Hex Driver - BioHorizons 0.050 Hex Driver; Regular; REF 134-350; BioHorizons Implant Systems; Inc.; Birmingham; AL 35243; non-sterile; Rx only. The product is part of a Prosthetic Instrumentation system Kit 51-K ; Manufactured by Innovision; Inc.; and the kits are assembled by BioHorizons. | 3 | Apr-13-2007 | BioHorizons Implant Systems Inc | Premature Wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear.
Z-0694-2007 | BioHorizons 0.050 Hex Driver; Long - BioHorizons 0.050 Hex Driver; Long; REF 300-351; BioHorizons Implant Systems; Inc.; Birmingham; AL 35243; non-sterile; Rx only | 3 | Apr-13-2007 | BioHorizons Implant Systems Inc | Premature Wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear.
Z-0693-2007 | BioHorizons 0.050 Hex Driver - BioHorizons 0.050 Hex Driver; Regular; REF 300-350; BioHorizons Implant Systems; Inc.; Birmingham; AL 35243; non-sterile; Rx only | 3 | Apr-13-2007 | BioHorizons Implant Systems Inc | Premature Wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear.
Z-0720-2007 | Urologix Targis - Urologix Microwave Delivery System 3.0 - 5.0. For use with prostatic urethra lengths of 3.0-5.0 cm only. Model TA1121D. | 2 | Apr-05-2007 | Urologix; Inc. | The catheter in the kit matches the description on the kit and pouch label; however; an incorrect label was placed onto the catheter during manufacturing.
Z-0721-2007 | Uroglogix Targis - Urologix Microwave Delivery System 2.5 - 3.5 For use with prostatic urethra lengths of 2.5-3.5 cm only. Model TA1321D. | 2 | Apr-05-2007 | Urologix; Inc. | The catheter in the kit matches the description on the kit and pouch label; however; an incorrect label was placed onto the catheter during manufacturing.
Z-0658-2007 | Terumo Perfusion System 8000 - Terumo Perfusion System 8000 Base; 5 pump; 220/240v; Model 8000; Catalog No. 16406. | 2 | Mar-27-2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.)
Z-0656-2007 | Terumo Perfusion System 8000 - Terumo Perfusion System 8000 Base; 5 pump; 115v; Model 8000; Catalog No. 16401. | 2 | Mar-27-2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.)
Z-0655-2007 | Terumo Perfusion System 8000 - Terumo Perfusion System 8000 Base; 4 pump; 115v; Model 8000; Catalog No. 16400. | 2 | Mar-27-2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.)
Z-0659-2007 | Terumo Perfusion System 8000 - Terumo Perfusion System 8000 Base; 5 pump; 100v; Model 8000; Catalog No. 16410. | 2 | Mar-27-2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.)
Z-0657-2007 | Terumo Perfusion System 8000 - Terumo Perfusion System 8000 Base; 4 pump; 220/240v; Model 8000; Catalog No. 16405. | 2 | Mar-27-2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.)
Z-0574-2007 | Abiomed - AB5000 Console Circulatory Support SystemCatalog Number: 0015-0000 | 2 | Mar-07-2007 | Abiomed; Inc. | Unit may alarm ''Low Flow'' and ''Low Pressure'' due to manufacturing material in the pressure source.
Z-0411-2007 | Rocker Reducer - Rocker Reducer; part of the SiLo Spinal System; REF 8880014; Rx only; Material: Stainless Steel; qty: 1 ea; non-sterile; Medtronic Sofamor Danek; 1800 Pyramid Place Memphis; TN 38132; USA; Manufactured at: Bartlett; Tennessee; USA | 2 | Feb-08-2007 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error; the proximal end may deform causing the protrusions that engage the associated implant to break off of the instrument during use.
Z-0300-2007 | BCI Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor - BCI 3404 Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor. Smiths Medical PM; Inc. This device is designed to provide full featured monitoring capabilities in a tabletop design. The system features an ECG cable interface; an SpO2 probe interface; display of patient and waveform data via an EL panel; power status LED; and the function keypad area consisting of six keys (on/off; waveform/trend; alarm silence; menu/enter; up arrow & down arrow). The monitor has a serial port that is used for data communications to a printer or computer and for analog outputs. | 2 | Dec-19-2006 | Smiths Medical PM; Inc. | An error occurred at the board manufacturing site which may affect the following parameters: - An electrical noise observed on the SpO2 pleth waveform generating and displaying erratic heart rates. % SpO2 reading displays dashes. - The ECG waveform may appear noisy; ECG readings appear accurate. -Respiration waveform may appear noisy and the respiration readings will be displayed as dashes.
Z-0243-2007 | BacT/ALERT - bioMerieux BacT/ALERT SN Anaerobic Culture Bottle-REF 259790 | 2 | Dec-05-2006 | bioMerieux; Inc. | Recovery Compromised-during manufacturing ambient air (including oxygen) was inadvertently introduced into the culture bottles; changing their reduction/oxidation potential thus leading to a compromised recovery of some obligate anaerobes and longer detection times.
Z-0190-2007 | ST360 Distal Thread Reduction Guide Pins - Zimmer Spine ST360 Distal Thread Reduction Guide Pins; Part number: 07.00684.00. Non-Sterile. Zimmer Spine; Inc.; 7375 Bush Lake Road; Minneapolis; MN 55439 U.S.A. The Distal Thread Reduction Guide Pin is a reusable instrument that is part of a set of instruments used to perform reduction surgeries with the ST360 Spinal Fixation System. The ST360 Spinal Fixation System is a temporary implant system used to correct spinal deformity and facilitate the biological process of spinal fusion. This system is intended for posterior use in the thoracic; lumbar and sacral areas of the spine. Implants of this system consist of hooks and/or screws connected to rods and are intended to be removed after solid fusion has occurred. This system includes polyaxial screws of varying diameters and lengths; fixed screws of varying diameters and lengths; rods in varying lengths; hooks in varying designs; fixed and adjustable transverse connectors. | 2 | Nov-16-2006 | Zimmer Spine; Inc. | Zimmer Spine; Inc. is conducting a RECALL of all manufacturing lots of ST360 Distal Thread Reduction Guide Pins due to an issue with the thread timing.
Z-0160-2007 | Fusion OMNI - Fusion OMNI ERCP Catheter; Wilson Cook Medical; GI Endoscopy; 4900 Bethania Station Road; Winston-Salem; NC 27105 | 2 | Nov-08-2006 | Cook Endoscopy | Injection through the flush port of these ERCP catheters may be compromised due to omission of a manufacturing activity.
Z-0155-2007 | Access Point Medical Heavy Duty Rollator with Curved Back - Access Point Medical-Heavy Duty Rollator with Curved Back (rolling walker); Product Code RL12011; with seat; curved backrest; 8'' wheels with locks; maximum capacity 400 lbs.;and wire basket. Manufactured for Access Point Medical; St. Louis; MO. Made in China | 2 | Nov-07-2006 | Access Point Medical LLC | The fork component on the wheel of the Rollator can break due to the use of incorrect manufacturing material and/or molding processes.
Z-0154-2007 | Access Point Medical Heavy Duty Rollator - Access Point Medical-Heavy Duty Rollator (rolling walker); Product Code RL12010HD; with seat; backrest; 8'' wheels with locks; maximum capacity 400 lbs.;and wire basket. Manufactured for Access Point Medical; St. Louis; MO. Made in China | 2 | Nov-07-2006 | Access Point Medical LLC | The fork component on the wheel of the Rollator can break due to the use of incorrect manufacturing material and/or molding processes.
Z-1491-06 | COBE Autotransfusion set. - Cobe STAT PAC Autotransfusion Set (made with Electa Wash Set); Cobe Product Codes: ASCBFE22; ASCBE22; AS9E55; AS9E22; AS9E12; AS4E55; AS4E22; AS4E17. Cobe Cardiovascular Inc. | 2 | Sep-09-2006 | Cobe Cardiovascular; Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets; possibly causing particles in the blood.
Z-1489-06 | dideco Surgical blood wash set - dideco Preassembled Surgical Wash Set; Electa Essential ; Cobe part numbers WS55E; WS125E; WS175E; WS225E; Cobe Cardiovascular; Inc. | 2 | Sep-09-2006 | Cobe Cardiovascular; Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets; possibly causing particles in the blood.
Z-1488-06 | dideco Surgical blood wash set - dideco Preassembled Surgical Wash Set; Compact ; Cobe part numbers WS55C; WS125C; WS225C; Cobe Cardiovascular; Inc. | 2 | Sep-09-2006 | Cobe Cardiovascular; Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets; possibly causing particles in the blood.
Z-1490-06 | COBE Autotransfusion set. - Cobe STAT PAC; Autotransfusion Set (Made with Compact Wash Sets); Cobe Product Codes: AS4C12; AS4C22; ASCBC12C; ASCBC22C; ASCBFC22C. Cobe Cardiovascular Inc. | 2 | Sep-09-2006 | Cobe Cardiovascular; Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets; possibly causing particles in the blood.
Z-1406-06 | Dri-tek D7120; BG-OX Cartridges Qualicheck2+ - D7120; BG-OX Cartridge; D7120 REF 946-003; and 946-001; cuvettes for the NPT7 series of blood gas analyzers. | 3 | Aug-26-2006 | Radiometer America Inc | QC Bias-a solution used in the manufacturing of the membrane had a different composition which introduces a bias of 20 mpH.
Z-1254-06 | VIASYS Healthcare Disposable Surface Electrodes - VIASYS Healthcare Disposable Bar Electrodes. Re-order No. 019-435600. Single Use only. Manufactured in the U.S.A. for VIASYS Healthcare; 5225 Verona Road; Bldg 2; Madison; WI USA 53711-4495..Each box contains 10 pouches of Bar Electrodes and each pouch contains 10 electrodes | 3 | Jul-25-2006 | Nicolet Biomedical Div of Viasys Healthcare | The Bar Electrodes do not meet performance specification for electrode impedance. A manufacturing defect was discovered in disposable bar electrodes; part number 019-435600; lot 604664; that has the potential to reduce the amplitude of a motor or sensory nerve action potential. These possible lower amplitudes will be measured only if the bar electrode is used as a bar.
Z-1234-06 | Nexcare High Performance Gauze Pad - 3M Nexcare First Aid High Performance Gauze Pad. Individually wrapped & sterile. catalog # 434-10. 3M Consumer Health Care; St. Paul; MN 55144-1000. Made in USA. UPC 05113166774. | 2 | Jul-13-2006 | 3M Company / Medical Division | After reviewing manufacturing records it has been determined that these lots inadvertently were not sterilized after packaging in individual boxes labeled as sterile product.
Z-0913-06 | Bard TruGuide Coxail Biopsy Needle with Depth Stop and additional Blunt Tip Stylet - Bard TruGuide Coxail Biopsy Needle with Depth Stop and additional Blunt Tip Stylet; Catalog Number C1816B; Device Listing Number E148687.The Bard TruGuide Coaxial Biopsy Needle is a three-part device consisting of an outer cannula with an attached female luer lock hub; and inner stylet with an attached male luer lock hub; and a flexible slip rink depth stop. | 2 | May-31-2006 | Bard Peripheral Vascular Inc | Manufacturing change that has introduced variability in the length of the inner stylet. When the inner stylet length is inserted into the cannula component; the exposed stylet length may be up to 8mm longer than expected.
Z-0890-06 | CONTAK RENEWAL 3 - CONTAK RENEWAL 3 CRT-D (model no. H170; H175; H177; H179). Cardiac Resynchronization Therapy Defibrillator. Sterilized with gaseous ethylene oxide. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. CONTAK RENEWAL 3 devices provide ventricular tachyarrhythmia and cardiac resynchronization therapies. | 3 | May-23-2006 | Guidant Corporation | Devices manufactured with batteries built on April 26 or 27; 2005; by a single operator; may have battery welds that may have been performed at settings outside of typical manufacturing specifications at the supplier.
Z-0891-06 | CONTAK RENEWAL 4 - CONTAK RENEWAL 4 CRT-D (model no. H190; H195; H199). Cardiac Resynchronization Therapy Defibrillator. Sterilized with gaseous ethylene oxide. Guidant Corporation; Cardiac Rhythm Management; 4100 Hamline Avenue North; St. Paul; MN 55112-5798 USA. CONTAK RENEWAL 4 devices provide ventricular tachyarrhythmia and cardiac resynchronization therapies. | 3 | May-23-2006 | Guidant Corporation | Devices manufactured with batteries built on April 26 or 27; 2005; by a single operator; may have battery welds that may have been performed at settings outside of typical manufacturing specifications at the supplier.
Z-0892-06 | VITALITY - VITALITY DS (model T125) and VITALITY 2 (models T165; 175) ICD (implantable cardiac defibrillator) | 3 | May-23-2006 | Guidant Corporation | Devices manufactured with batteries built on April 26 or 27; 2005; by a single operator; may have battery welds that may have been performed at settings outside of typical manufacturing specifications at the supplier.
Z-0830-06 | Zinetics Reference pH Catheter - Medtronic Zinetics 24ME Multi-Use External Reference pH Catheter. For Gastroesophageal pH Measurements. Catalog number 9012P2121. Distributed by Medtronic; Inc.; 710 Medtronic Parkway NE; Minneapolis; MN 55432-5604; USA. | 2 | May-03-2006 | Medtronic Gastroenterology / Urology | Lot 0001491 has a manufacturing defect in the sensor spacing. Sensors have been placed 5 cm apart instead of 15 cm apart. Medtronic has determined that patient safety is not compromised. Clinicians however should review and reassess the results of any patient studies for which catheters from lot 0001491 were used.
Z-0743-06 | METRx II Tube - METRx II Tube; 20mm x 6cm stainless steel tubular retractor; Part No. 9569723 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0725-06 | METRx II Tube - METRx II Tube; 14mm x 9cm stainless steel tubular retractor; Part No. 9569666 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a Manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0727-06 | METRx II Tube - METRx II Tube; 16mm x 4cm stainless steel tubular retractor; Part No. 9569671 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0751-06 | METRx II Tube - METRx II Tube; 22mm x 7cm stainless steel tubular retractor; Part No. 9569734 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0753-06 | METRx II Tube - METRx II Tube; 22mm x 9cm stainless steel tubular retractor; Part No. 9569736 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0742-06 | METRx II Tube - METRx II Tube; 20mm x 5cm stainless steel tubular retractor; Part No. 9569722 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0744-06 | METRx II Tube - METRx II Tube; 20mm x 7cm stainless steel tubular retractor; Part No. 9569724 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0747-06 | METRx II Tube - METRx II Tube; 22mm x 3cm stainless steel tubular retractor; Part No. 9569730 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0741-06 | METRx II Tube - METRx II Tube; 20mm x 4cm stainless steel tubular retractor; Part No. 9569721 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0734-06 | METRx II Tube - METRx II Tube; 18mm x 4cm stainless steel tubular retractor; Part No. 9569681 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0750-06 | METRx II Tube - METRx II Tube; 22mm x 6cm stainless steel tubular retractor; Part No. 9569733 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0723-06 | METRx II Tube - METRx II Tube; 14mm x 7cm stainless steel tubular retractor; Part No. 9569664 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0748-06 | METRx II Tube - METRx II Tube; 22mm x 4cm stainless steel tubular retractor; Part No. 9569731 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0722-06 | METRx II Tube - METRx II Tube; 14mm x 6cm stainless steel tubular retractor; Part No. 9569663 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0732-06 | METRx II Tube - METRx II Tube; 16mm x 9cm stainless steel tubular retractor; Part No. 9569676 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0724-06 | METRx II Tube - METRx II Tube; 14mm x 8cm stainless steel tubular retractor; Part No. 9569665 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0749-06 | METRx II Tube - METRx II Tube; 22mm x 5cm stainless steel tubular retractor; Part No. 9569732 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0746-06 | METRx II Tube - METRx II Tube; 20mm x 9cm stainless steel tubular retractor; Part No. 9569726 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0738-06 | METRx II Tube - METRx II Tube; 18mm x 8cm stainless steel tubular retractor; Part No. 9569685 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0733-06 | METRx II Tube - METRx II Tube; 18mm x 3cm stainless steel tubular retractor; Part No. 9569680 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0752-06 | METRx II Tube - METRx II Tube; 22mm x 8cm stainless steel tubular retractor; Part No. 9569735 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0731-06 | METRx II Tube - METRx II Tube; 16mm x 8cm stainless steel tubular retractor; Part No. 9569675 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0735-06 | METRx II Tube - METRx II Tube; 18mm x 5cm stainless steel tubular retractor; Part No. 9569682 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0720-06 | METRx II Tube - METRx II Tube; 14mm x 4cm stainless steel tubular retractor; Part No. 9569661 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0730-06 | METRx II Tube - METRx II Tube; 16mm x 7cm stainless steel tubular retractor; Part No. 9569674 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0745-06 | METRx II Tube - METRx II Tube; 20mm x 8cm stainless steel tubular retractor; Part No. 9569725 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0740-06 | METRx II Tube - METRx II Tube; 20mm x 3cm stainless steel tubular retractor; Part No. 9569720 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0721-06 | METRx II Tube - METRx II Tube; 14mm x 5cm stainless steel tubular retractor; Part No. 9569662 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0719-06 | METRx II Tube - METRx II Tube; 14mm x 3cm stainless steel tubular retractor; Part No. 9569660 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error; after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0726-06 | METRx II Tube - METRx II Tube; 16mm x 3cm stainless steel tubular retractor; Part No. 9569670 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving; the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0737-06 | METRx II Tube - METRx II Tube; 18mm x 7cm stainless steel tubular retractor; Part No. 9569684 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0729-06 | METRx II Tube - METRx II Tube; 16mm x 6cm stainless steel tubular retractor; Part No. 9569673 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0728-06 | METRx II Tube - METRx II Tube; 16mm x 5cm stainless steel tubular retractor; Part No. 9569672 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0739-06 | METRx II Tube - METRx II Tube; 18mm x 9cm stainless steel tubular retractor; Part No. 9569686 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0736-06 | METRx II Tube - METRx II Tube; 18mm x 6cm stainless steel tubular retractor; Part No. 9569683 | 2 | Apr-12-2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error after autoclaving the Tubular Retractor coating; Physical Vapor Deposition (PVD) black Titanium Nitride (TiN); may rub off.
Z-0695-06 | - Fisher & Paykel Healthcare Adult Respiratory Ventilator Circuit-Dual Heated Catalogue Number RT111 | 2 | Mar-29-2006 | Fisher & Paykel Healthcare Inc | Risk of fire. The manufacturing process may have the potential to damage the insulation of the heater wire in circuits within the specified lot range. This may lead to a potential risk of fire.
Z-0696-06 | - Fisher & Paykel Healthcare Adult Respiratory Ventilator Circuit-Dual Heated Catalogue Number RT210 | 2 | Mar-29-2006 | Fisher & Paykel Healthcare Inc | Risk of fire. The manufacturing process may have the potential to damage the insulation of the heater wire in circuits within the specified lot range. This may lead to a potential risk of fire.
Z-0697-06 | - Fisher & Paykel Healthcare Adult Respiratory Ventilator Circuit-Dual Heated Catalogue Number RT211 | 2 | Mar-29-2006 | Fisher & Paykel Healthcare Inc | Risk of fire. The manufacturing process may have the potential to damage the insulation of the heater wire in circuits within the specified lot range. This may lead to a potential risk of fire.
Z-0694-06 | - Fisher & Paykel Healthcare Adult Respiratory Ventilator Circuit-Dual Heated Catalogue Number RT110 | 2 | Mar-29-2006 | Fisher & Paykel Healthcare Inc | Risk of fire. The manufacturing process may have the potential to damage the insulation of the heater wire in circuits within the specified lot range. This may lead to a potential risk of fire.
Z-0612-06 | Baxter Auto Syringe AS50 Infusion Pump - Baxter Auto Syringe AS50 Infusion Pump; product code 1M8550; Baxter Healthcare Corporation; Medication Delivery Division; Deerfield; IL 60015 U.S.A. | 2 | Mar-16-2006 | Baxter Healthcare Corp. | A non-conforming gasket was used in manufacturing of the specific pumps; causing the pumps to be more susceptible to electro-static discharge; which could increase the potential for the pump to fail during use and alarm.
Z-0504-06 | Hospira Latex Free Blood Secondary Set - Hospira Latex Free Blood Secondary Set; 26 inch with Secure Lock; Made in Costa Rica; Hospira; Inc.; Lake Forest; IL 60045 U.S.A. | 2 | Feb-08-2006 | Hospira Inc. | The I.V. administrations sets were stolen in-transit between the manufacturing site and the sterilizer; resulting in non-sterilized product possibly being on the market.
Z-0506-06 | Hospira Transpac IV Monitoring Kit - Hospira Transpac IV Monitoring Kit; 60 inch disposable transducer monitoring kit with second stopcock by transducer; continuous flush device and drip chamber; Made in Costa Rica; Hospira; Inc.; Lake Forest; IL 60045 U.S.A. | 2 | Feb-08-2006 | Hospira Inc. | The I.V. administrations sets were stolen in-transit between the manufacturing site and the sterilizer; resulting in non-sterilized product possibly being on the market.
Z-0505-06 | LifeShield Latex-Free Primary I.V. Set - LifeShield Latex-Free Primary I.V. Set; Convertible Pin; 83 inch; with Backcheck Valve and 2 Clave Ports; Piggyback with Option-Lok; Made in Costa Rica; Hospira; Inc.; Lake Forest; IL 60045 U.S.A. | 2 | Feb-08-2006 | Hospira Inc. | The I.V. administrations sets were stolen in-transit between the manufacturing site and the sterilizer; resulting in non-sterilized product possibly being on the market.
Z-0392-06 | Medtronic Slimline Catheter - Medtronic Slimline Single-Use Internal Reference pH Catheter for Gastroesophageal pH measurements. Models 9012P3101; 9012P3111; 9012P3121; 9012P3131 and 9012P3141. Distributed by / Medtronic Inc. 710 Medtronic Parkway NE; Minneapolis; MN 55432-5604 USA | 3 | Jan-13-2006 | Medtronic Gastroenterology / Urology | Certain lots have a manufacturing defect that allows the water pushed through the catheter to travel back into the connector of the Digitrapper pH recorder. Medtronic has determined that patient safety is not compromised; but catheter function may be affected and damage may result to the Digitrapper pH recorder.
Z-0354-06 | Respironics PLV Continuum - The PL V Continuum ventilator is a microprocessor controlled; compressorbased; mechanical ventilator. It is intended to control or assist breathing by delivering room air to the patient. PL V Continuum utilizes an internal compressor to generate compressed air for delivery to the patient. Breath delivery is controlled by software algorithms. The user interface on PL V Continuum has a membrane keypad with indicator Light Emitting Diodes (LED) for the selection and acceptance of patient settings and for the display of alarm conditions. PL V Continuum is capable of providing the following types of ventilatory support:. Positive Pressure Ventilation; delivered either invasively (via endotracheal or tracheostomy tube) or non-invasively (via mask or mouthpiece).. Assist/Control; Spontaneous Intermittent Mandatory Ventilation (SIMV) or Continuous Positive Airway Pressure (CP AP) modes of ventilation.. Volume-Controlled (VC). Available in AlC and SIMV.. Pressure-Controlled (PC). Available in AlC and SIMV.. Pressure Support (PS). Available in SIMV and SPONT. | 2 | Jan-06-2006 | Respironics California Inc | During product testing of the PLV Continuum Ventilator's Safety Valve/Pressure Relief Valve (SV/PRV) at our manufacturing facility; Respironics found that when the PLY Continuum Ventilator is operated in the 'Face Up' orientation; the SV /PRV may stick in the open position; which may result in loss of delivered gas volume to the patient.
Z-0266-06 | Symphony DR and ELA Rhapsody DR - Symphony DR model 2550 and ELA Rhapsody DR model 2510. A dual chamber rate-responsive pacemaker that uses both minute ventilation and accelerometer sensors. Sterilization with Ethylene Oxide Gas. Sorin Group; ELA medical; 98; rue Maurice Amoux; 92120 Montrouge - France. Made in France. | 2 | Dec-02-2005 | ela Medical Llc | Through post-market surveillance; ELA Medical has observed that a no-output condition could occur in a limited number of Symphony or rhapsody pacemakers. The no-output condition could occur due to metal migration caused by a specific manufacturing process.
Z-0267-06 | ELA Rhapsody SR - ELA Rhapsody SR model 2210. A single chamber rate-responsive pacemaker that uses both minute ventilation and accelerometer sensors. Sterilization with Ethylene Oxide Gas. Sorin Group; ELA medical; 98; rue Maurice Amoux; 92120 Montrouge - France. Made in France. | 2 | Dec-02-2005 | ela Medical Llc | Through post-market surveillance; ELA Medical has observed that a no-output condition could occur in a limited number of Symphony or rhapsody pacemakers. The no-output condition could occur due to metal migration caused by a specific manufacturing process.
Z-0144-06 | D-Stat Radial Hemostat Band - D-Stat 2 Dry Hemostatic Bandage; model number 3010. Sterile; latex free single use device. Vascular Solutions; Inc. 6464 Sycamore Court; Minneapolis; MN 55369 USA. Each D-Stat 2 Dry hemostatic bandage (D-Stat 2 Dry) consists of the following components: -One (1) scored lyophilized pad consisting of thrombin; sodium carboxylmethyulcellulose and calcium chloride; -Two (2) Adhesive bandages. | 2 | Nov-16-2005 | Vascular Solutions; Inc. | A defect affecting the seals of the paper pouch holding the adhesive bandage provided with the D-Stat products was noted during manufacturing. The integrity of the seals of the D-stat Dry pads which are packaged separately in a foil pouch is not in question.
Z-0143-06 | D-Stat Radial Hemostat Band - D-Stat Hemostat Band; model number 3500. Sterile; latex free single use device. Vascular Solutions; Inc. 6464 Sycamore Court; Minneapolis; MN 55369 USA. Each D-Stat Radial hemostatic band (D-Stat Radial) consists of the following components: -Lyophilized pad consisting of thrombin; sodium carboxylmethyulcellulose and calcium chloride in a nonwoven gauze; -Application device consisting of an adjustable retention strap; collar and attached gauze pad. | 2 | Nov-16-2005 | Vascular Solutions; Inc. | A defect affecting the seals of the paper pouch holding the adhesive bandage provided with the D-Stat products was noted during manufacturing. The integrity of the seals of the D-stat Dry pads which are packaged separately in a foil pouch is not in question.
Z-0142-06 | D-Stat Dry Hemostatic adhesive bandage - D-Stat Dry Hemostatic Bandage; model number 3000. Sterile; latex free single use device. Vascular Solutions; Inc. 6464 Sycamore Court; Minneapolis; MN 55369 USA. Each D-Stat Dry hemostatic bandage (D-Stat Dry) consists of the following components: -Lyophilized pad consisting of thrombin; sodium carboxylmethyulcellulose and calcium chloride in a nonwoven gauze; -Adhesive bandage | 2 | Nov-16-2005 | Vascular Solutions; Inc. | A defect affecting the seals of the paper pouch holding the adhesive bandage provided with the D-Stat products was noted during manufacturing. The integrity of the seals of the D-stat Dry pads which are packaged separately in a foil pouch is not in question.
Z-1585-05 | Bard Luminexx 3 Biliary Stent and Delivery System - Bard Luminexx 3 Biliary Stent and Delivery System | 2 | Sep-24-2005 | Bard Peripheral Vascular Inc | Angiomed Material Review Board reviewing the manufacturing process observed an increase in the internal scrap rate (failure mode: broken stent struts) of nitinol stents and determined the root cause to be a manufacturing change from polymer to aluminum fixtures.
Z-1113-05 | Healthwear - Roche Diagnostic HealthWear brand Professional System for weight management; Catalog # 04344456001. | 3 | Aug-12-2005 | Roche Diagnostics Corp. | Lack of adequate manufacturing controls.
Z-1085-05 | - Puritan Bennett 840 Series Ventilator Power SupplyModel/Catalog Numbers: 4-076314-00; 4-076314-SP | 2 | Aug-06-2005 | Puritan Bennett | Manufacturing defect at contract manufacturer; employee error; poor training; improper inspection.
Z-1068-05 | Sequoia - Sequoia Oxymatic electronic oxygen conservers. Models OM-302; OM-311 ***and OM-312***. | 2 | Aug-03-2005 | Chad Therapeutics Inc | Shot size below manufacturing specifications. Risk of less than the specified quantity of oxygen delivered per pulse.
Z-0852-05 | Continuous Nerve Block Needle - Arrow Continuous Nerve Block Needle; Latex-Free* Rx only*Sterile*Arrow International Inc.; 2400 Bernville; Road; Reading; PA 19605 USA | 2 | Jun-01-2005 | Arrow International Inc | Possible difficulty in passing the catheter through the needle due to a reduced ID of the needle caused by over crimping of the needle during the manufacturing process.
Z-0434-05 | Guiding Catheter - VISTA BRITE TIP RDC(1) Guiding Catheter. Product Code 67021055; lot X0704002; Precutaneous Catheter. | 3 | Jan-26-2005 | Cordis Corporation | Two complaints have been received and confirmed; regarding a lot; indicating that the distal end had an 'RDC' shape rather than the 'RDC(1)' shape. The RDC shape has a larger curve. The product was manufactured with the incorrect forming wire during the guiding catheter manufacturing process.
Z-0393-05 | - Weck; Endo5; Hem-o-lok¿ ML; Ref 544965; Hem-o-lok ML Endoscopic Clip Applier; Non-Sterile Clean and sterilize before use. Weck Closure Systems¿; TFX Medical Ltd.; Halifax Road; P.O. Box 138; High Wycombe HP123NB U.K. | 3 | Jan-13-2005 | Weck | The flush port on some appliers was occluded with an epoxy resin during manufacturing.
Z-1341-04 | ARCHITECT - ARCHITECT c8000 Processing Module; List Number 01G06-01. Manufactured by Toshiba Corporation; Medical Systems Division; Japan. Distributed by Abbott Laboratories Diagnostic Division; Irving; Texas | 3 | Aug-13-2004 | Abbott Laboratories; Inc | Process change in the manufacturing of the AC/DC driver board used on the analyzer may result in one or more of the solid state relays to fail during use.
Z-0992-04 | ABX Diagnostics and Horiba - ABX Diagnostics Minotrol 16 Hematology Controls; Normal and High; in vials distributed by ABX Diagnostics; Irvine; CA; andHoriba LC-TROL 16 (Hematology Reference Controls); Normal and High; in vials distributed by Horiba Ltd.; Kyoto; Japan | 3 | Jul-20-2004 | R & D Systems;Inc | Clots in hematology controls resulted in low red blood cell counts when used in blood analyzers because control values were outside the expected range due to the clotting which took place after manufacturing.
Z-1066-04 | Visulas 532 s with Visulink 532/U - Surgical powered laser instrument; used for photocoagulation of ocular tissues in the treatment of diseases of the eye. | 1 | Jul-20-2004 | Carl Zeiss Ophthalmic Systems Inc | The coating of laser mirrors in a defined manufacturing lot of the Visulink 532/U devices may partially tear open and loosen from the base plate. The mirror may misdirect the laser beam to an unintended target in or on the eye. Retinal bleeding and burns may result.
Z-0722-04 | - Esprit Ventilator. | 2 | Jul-20-2004 | Respironics California Inc | Valves may stick in the closed position due to manufacturing error.
Z-0536-04 | Four Shooter Saeed Multi-Band Ligator - Four Shooter Saeed Multi-Band Ligator | 3 | Jul-20-2004 | Wilson-Cook Medical Inc | Product label lacks the symbols specifying the meaning of the manufacturing and expiration date.
Z-0506-04 | Gas analyzer - AG-920RA Multi-gas unit analyzer | 2 | Jan-31-2004 | Nihon Kohden America Inc | Manufacturing error results in screw affecting electromagnetic interferences. Also metal lid is not tight enough.
Z-0658-03 | Flexgard; off-the-shelf knee brace - Flexgard; off-the-shelf knee brace; product number 14761106 | 2 | Mar-19-2003 | De Royal Industries | Manufacturing defect possibly causing screws to come loose causing the brace to collapse during use.
Z-0630-03 | OneTouch SureStep - OneTouch SureStep Glucose Test Strips (consumer use); OneTouch SureStep Pro Glucose Test Strips (hospital use) | 2 | Mar-12-2003 | Lifescan Inc | The Test Strips have manufacturing errors; in that the 'Confirmation Dot' is off-center or split in half; and defective bottle Cap.

## Sheet1
Year | #of Recalls | #Mfg Issues
2003 | 2 | 2
2004 | 6 | 6
2005 | 12 | 9
2006 | 63 | 18
2007 | 28 | 10
2008 | 23 | 10
2009 | 12 | 11
2010 | 158 | 31
2011 | 39 | 16
2012 | 35 | 20